IVI-3.15: IviLxiSync Specification August 23, 2018 Revision 2.0

Important Information

The IviLxiSync Specification (IVI-3.15) is authored by the IVI Foundation member companies. For a vendor membership roster list, please visit the IVI Foundation web site at www.ivifoundation.org.

The IVI Foundation wants to receive your comments on this specification. You can contact the Foundation through the web site at www.ivifoundation.org.

- Warranty

The IVI Foundation and its member companies make no warranty of any kind with regard to this material, including, but not limited to, the implied warranties of merchantability and fitness for a particular purpose. The IVI Foundation and its member companies shall not be liable for errors contained herein or for incidental or consequential damages in connection with the furnishing, performance, or use of this material.

- Trademarks

Product and company names listed are trademarks or trade names of their respective companies.

No investigation has been made of common-law trademark rights in any work.

IviLxiSync Interface Specification	7

1. Overview of the IviLxiSync Interface Specification	9

1.1 Introduction	9

1.2 IviLxiSync Interface Overview	9

1.3 References	9

1.4 Definitions of Terms and Acronyms	10

1.5 IviLxiSync Device Model	10

1.5.1 High-Level Device Model	10

1.6 Integrating IviLxiSync With Existing Classes	10

1.7 Implementing IviLxiSync for an IVI-C Instrument Driver	11

1.8 Implementing IviLxiSync for an IVI-COM Instrument Driver	11

2. IviLxiSync Repeated Capabilities	12

2.1 Repeated Capability Names	12

2.1.1 IviLxiSyncArmSource	12

2.1.2 IviLxiSyncArmAlarm	12

2.1.3 IviLxiSyncTriggerAlarm	12

2.1.4 IviLxiSyncTriggerSource	13

2.1.5 IviLxiSyncEvent	13

2.1.6 Reserved Repeated Capability Identifiers	13

2.1.7 Custom Repeated Capability Identifiers	14

2.1.8 Repeated Capability Identifier Case Sensitivity	14

2.1.9 Repeated Capability Implementation Requirements	15

2.2 IviLxiSync Group Names	15

2.2.1 IviLxiSync Group Names	15

2.3 Boolean Attribute and Parameter Values	17

2.4 .NET Namespace	17

3. IviLxiSyncArm Subsystem	18

3.1 Behavior Model	18

3.2 IviLxiSyncArm Attributes	21

3.2.1 Arm Count	22

3.2.2 Arm Alarm Count	23

3.2.3 Arm Alarm Enabled	24

3.2.4 Arm Alarm Item (IVI-COM and IVI.NET Only)	25

3.2.5 Arm Alarm Name (IVI-COM and IVI.NET Only)	26

3.2.6 Arm Alarm Period	27

3.2.7 Arm Alarm Repeat Count	28

3.2.8 Arm Alarm Time (IVI.NET Only)	29

3.2.9 Arm Alarm Time Seconds (IVI-C and IVI-COM Only)	30

3.2.10 Arm Alarm Time Fraction (IVI-C and IVI-COM Only)	31

3.2.11 Arm Delay	32

3.2.12 Arm Source Count	33

3.2.13 Arm Source Detection	34

3.2.14 Arm Source Enabled	36

3.2.15 Arm Source EventId	37

3.2.16 Arm Source Filter	38

3.2.17 Arm Source Item (IVI-COM and IVI.NET Only)	40

3.2.18 Arm Source Name (IVI-COM and IVI.NET Only)	41

3.2.19 Arm Source Or Enabled	42

3.3 IviLxiSyncArm Functions	43

3.3.1 Add Arm Alarm	44

3.3.2 Add Arm Source	46

3.3.3 Configure Arm Alarm	48

3.3.4 Configure Arm Source	50

3.3.5 Disable All Arm Alarms	51

3.3.6 Disable All Arm Sources	52

3.3.7 Get Arm Alarm Name (IVI-C Only)	53

3.3.8 Get Arm Source Name (IVI-C Only)	54

3.3.9 Remove Arm Alarm	55

3.3.10 Remove Arm Source	56

3.3.11 Remove All Custom Arm Alarms	57

3.3.12 Remove All Custom Arm Sources	58

4. IviLxiSyncTrigger Subsystem	59

4.1 Behavior Model	59

4.2 IviLxiSyncTrigger Attributes	60

4.2.1 Trigger Alarm Count	61

4.2.2 Trigger Alarm Enabled	62

4.2.3 Trigger Alarm Item (IVI-COM and IVI.NET Only)	63

4.2.4 Trigger Alarm Name (IVI-COM and IVI.NET Only)	64

4.2.5 Trigger Alarm Period	65

4.2.6 Trigger Alarm Repeat Count	66

4.2.7 Trigger Alarm Time (IVI.NET Only)	67

4.2.8 Trigger Alarm Time Seconds (IVI-C and IVI-COM Only)	68

4.2.9 Trigger Alarm Time Fraction (IVI-C and IVI-COM Only)	69

4.2.10 Trigger Count	70

4.2.11 Trigger Source	71

4.2.12 Trigger Source Count	72

4.2.13 Trigger Source Delay	73

4.2.14 Trigger Source Detection	74

4.2.15 Trigger Source EventId	75

4.2.16 Trigger Source Item (IVI-COM and IVI.NET Only)	76

4.2.17 Trigger Source Name (IVI-COM and IVI.NET Only)	77

4.2.18 Trigger Source Filter	78

4.3 IviLxiSyncTrigger Functions	80

4.3.1 Add Trigger Alarm	81

4.3.2 Add Trigger Source	83

4.3.3 Configure Trigger Alarm	85

4.3.4 Configure Trigger Source	87

4.3.5 Disable All Trigger Alarms	88

4.3.6 Get Trigger Alarm Name (IVI-C Only)	89

4.3.7 Get Trigger Source Name (IVI-C Only)	90

4.3.8 Remove Trigger Alarm	91

4.3.9 Remove Trigger Source	92

4.3.10 Remove All Custom Trigger Sources	93

4.3.11 Remove All Trigger Alarms	94

5. IviLxiSyncEvent Subsystem	95

5.1 Behavior Model	95

5.2 IviLxiSyncEvent Attributes	96

5.2.1 Event Count	97

5.2.2 Event Destination Path	98

5.2.3 Event Drive Mode	101

5.2.4 Event Item (IVI-COM and IVI.NET Only)	103

5.2.5 Event Name (IVI-COM  and IVI.NET Only)	104

5.2.6 Event Slope	105

5.2.7 Event Source	106

5.2.8 Event Wired OR Bias Mode	108

5.3 IviLxiSyncEvent Functions	110

5.3.1 Add Event	111

5.3.2 Configure Event	113

5.3.3 Disable All Events	115

5.3.4 Get Event Name (IVI-C Only)	116

5.3.5 Remove Event	117

5.3.6 Remove All Custom Events	118

6. IviLxiSyncEventLog Subsystem	119

6.1 IviLxiSyncEventLog Attributes	119

6.1.1 Event Log Entry Count	120

6.1.2 Event Log Enabled	121

6.2 IviLxiSyncEventLog Functions	122

6.2.1 Clear Event Log Entries	123

6.2.2 Get Next Event Log Entry	124

7. IviLxiSyncTime Subsystem	125

7.1 IviLxiSyncTime Attributes	125

7.1.1 Is Time Master	126

7.1.2 Is Time Synchronized	127

7.1.3 System Time (IVI.NET Only)	128

7.2 IviLxiSyncTime Functions	129

7.2.1 Get System Time (IVI-C and IVI-COM Only)	130

8. Attribute ID Definitions	131

9. Attribute Value Definitions	133

10. Function Parameter Value Definitions	135

11. Error and Completion Code Value Definitions	136

11.1 IVI.NET IviLxiSync Exceptions and Warnings	138

11.1.1 AlarmDoesNotExistException	139

11.1.2 AlarmExistsException	140

11.1.3 AlarmTimeInvalidException	141

11.1.4 CannotResolveReservedRepeatedCapabilityException	142

11.1.5 EventSourceDoesNotExistException	143

11.1.6 EventSourceExistsException	144

11.1.7 EventSourceNotSetException	145

11.1.8 InvalidEventSourceException	146

11.1.9 OutOfEventResourcesException	147

11.1.10 WiredOrModeInvalidException	148

12. Hierarchies	149

12.1 .NET Hierarchy	149

12.1.1 IviLxiSync .NET Interfaces	151

12.1.2 Interface Reference Properties	152

12.2 COM Hierarchy	152

12.2.1 IviLxiSync COM Interfaces	154

12.2.2 COM Interfaces	155

12.2.3 COM Interface Reference Properties	156

12.2.4 COM Category	156

12.2.5 COM Interface Accessibility	157

12.3 C Function Hierarchy	157

12.4 C Attribute Hierarchy	158

IviLxiSync Interface Specification

- Revision History

| Table 1. Specification Revisions |  |  |
| --- | --- | --- |
| Revision Number | Date of Revision | Revision Notes |
| Revision 1.0 | April 4, 2006 | First approved version. Initial Revision |
|  | September 20, 2006 | Editorial revision to 1.0 3.1 Clarify use of ARM is optional and recommend that no ARM config needed if not used 2.1.3, 2.1.4, 4.2.10 Clarify namespace overlap between TriggerAlarm, TriggerSource 2.1.1, 2.1.2 Clarify namespace between ArmAlarm and ArmSource 2.1.6 Pointed out that LXI reserves identifiers beginning with “LXI” 2.1.9 Pointed out that Custom capability groups are optional 4.1 Pointed out that trigger defaults should allow Initiate to start measurement 3.2.15, 4.2.17 Corrected prose regarding filter related to implied monitoring of send port. Eliminated syntactically incorrect example. 5.2.2 Pointed out that LXI reserves identifiers beginning with “LXI” 5.2.7 Corrected table and pointed out applicability beyond LXI 3.2.6, 3.2.7, 4.2.6, 4.2.7 Clarification of the behavior of Arm/Trigger count and interaction with Arm/Trigger period. |
|  | June 19, 2008 | Editorial Revision: Update the IVI Foundation contact information in the Important Information section to remove obsolete address information and refer only to the IVI Foundation web site. |
| Revision 2.0 | June 9, 2010 | Incorporated .NET |
| Revision 2.0 | October 7, 2014 | Editorial Change: Change the COM defined values listed in sections 3.2.13 and 4.2.14 to match the names in section 9 and the type library. |
| Revision 2.0 | August 23, 2018 | Editorial Change: Errant comments about case preserving and case insensitive were removed. |

- API Versions

| Architecture | Drivers that comply with version 2.0 comply with all of the versions below |
| --- | --- |
| C | 1.0, 2.0 |
| COM | 1.0, 2.0 |
| .NET | 1.0, 2.0 |

Drivers that comply with this version of the specification also comply with earlier, compatible, versions of the specification as shown in the table above.  The driver may benefit by advertising that it supports all the API versions listed in the table above.

# Overview of the IviLxiSync Interface Specification

## Introduction

This section introduces the IviLxiSync Interface Specification. This section summarizes the IviLxiSync Interface Specification  itself and contains general information that the reader may need in order to understand, interpret, and implement aspects of this specification. These aspects include the following:

- IviLxiSync API Overview

- References

- The definitions of terms and acronyms

## IviLxiSync Interface Overview

This specification defines the API for controlling the arming, triggering, and event functionality of LXI devices.  The functionality defined in the IviLxiSync Interface Specification is pertinent to LXI Class A and Class B devices, but is independent of the IVI instrument class supported by the device, if any.  The IviLxiSync Interface Specification conceptualizes an LXI device as an instrument that can “listen” for LXI trigger bus events or LXI LAN-based events and can fire such events.  Inbound events control the device’s arming and triggering subsystem for performing measurements and other operations.  A device may include conventional inbound events in addition to the LXI events.  Outbound events can be used to notify other LXI devices of specific conditions.

The IviLxiSync API is divided into the following five subsystems:

- IviLxiSyncArm

- IviLxiSyncTrigger

- IviLxiSyncEvent

- IviLxiSyncEventLog

- IviLxiSyncTime

- The IviLxiSyncArm subsystem controls when triggers are accepted.

- This IviLxiSyncTrigger subsystem controls when the LXI device triggers a measurement or other operation.

- The IviLxiSyncEvent subsystem controls when the LXI device signals specific conditions to other LXI devices.

- The IviLxiSyncEventLog subsystem provides access to the event logging features of an LXI device.

- The IviLxiSyncTime subsystem exposes functionality for access to the 1588 timebase of the LXI bus.

## References

The following documents and specifications are related to this specification:

- IVI-3.1: Driver Architecture Specification

- IVI-3.2: Inherent Capabilities Specification

- IVI-3.4: API Style Guide

- IVI-3.18: IVI.NET Utility Classes and Interfaces Specification

- IVI-5.0: Glossary

- LXI Standard

## Definitions of Terms and Acronyms

Refer to IVI-5: Glossary and to LXI Standards Definition: Appendix A – Glossary of Terms for a description of the terms and acronyms used in this specification. This specification does not define any additional terms.

## IviLxiSync Device Model

This section describes the logical model that is controlled by the IviLxiSync API.

### High-Level Device Model

The diagram below shows a high-level view of the major components of an LXI device from a trigger and synchronization perspective.  The three main IviLxiSync subsystems controlled by the interface in this specification (IviLxiSyncArm, IviLxiSyncTrigger, and IviLxiSyncEvent) are shown in the diagram as the “Arm Logic”, “Trigger Logic”, and “Event Logic” blocks, respectively.

The Arm Logic and Trigger Logic blocks receive input from both the LXI trigger bus and the LAN input registers.  LXI-defined LAN events (LAN0..7), as well as custom LAN events can drive the arm-trigger logic.  The output of the arm-trigger logic drives the Arm-Trigger State Machine, which is detailed further in the following section.

The Event Logic controls signals that are emitted from the LXI device.  It specifies which LXI devices are notified and what type of event notification is used.  The Event Logic can drive any of the LXI trigger bus lines, any LXI-defined LAN event (LAN0..7) or a custom LAN event.

Figure 1.  High-Level LXI Device Model

## Integrating IviLxiSync With Existing Classes

The IviLxiSync API is used with some other instrument API that supplies the measurement or stimulus capabilities. This existing API could be an IVI Class or it could be a purely custom API.  Usually, that other API will already define a trigger source attribute.  That attribute could be either an enumerated value or a string.  Since the IviLxiSync API also includes a trigger source, the two trigger source attributes and corresponding queries need to be integrated.

To integrate the IviLxiSync API with existing API’s, the following approach should be taken:

- The instrument should configure the trigger source to the value most recently configured by the customer, with the corresponding behavior model.  That is, if the customer configures the IviLxiSync trigger source via the IviLxiSync API, the IviLxiSync API will set the source and govern the behavior model.  If a class driver or specific driver API is used to set the trigger source (for instance, the DMM trigger source attribute), then the source specified via that API, and the corresponding behaviors, become active.

- When querying the trigger source on a given API, if the trigger source was set via that API, then the corresponding value is returned.  If the value was set via the alternate API and no corresponding value is available, then the driver should generate an error.

- The IviLxiSync API sets the trigger source using a string. Instruments are therefore, free to support additional trigger sources through that API.

## Implementing IviLxiSync for an IVI-C Instrument Driver

The IviLxiSync API is used in conjunction with existing IVI Foundation-defined API’s for instrument drivers. The IviLxiSync API defines some C interface elements which, if standard IVI naming conventions were followed, would conflict with existing IVI Foundation Class and Inherent interfaces for IVI-C instrument drivers. When implementing IviLxiSync for an IVI-C driver, different naming guidelines must therefore be followed for the IviLxiSync interface elements:

If you are implementing the IviLxiSync API as part of an IVI-C specific driver’s API:

- Replace “IviLxiSync_” in method names with “<Prefix>_IviLxiSync”, where <Prefix> is the actual driver prefix used for all specific driver function names.

Replace “IVILXISYNC_ATTR“ in attribute identifiers with “<PREFIX>_ATTR_IVILXISYNC”, where <PREFIX> is the  actual driver prefix using uppercase characters.

Replace “IVILXISYNC_ERROR” in error code identifiers with “<PREFIX>_ERROR_IVILXISYNC”, where <PREFIX> is the actual driver prefix using uppercase characters.

Replace “IVILXISYNC_VAL” in value identifiers  with “<PREFIX>_VAL_IVILXISYNC”, where <PREFIX> is the actual driver prefix using uppercase characters.

If you are implementing the IviLxiSync API as a common library to be shared between multiple specific driver API’s:

Use the attribute, value, function, and error code identifiers with the “IviLxiSync” or “IVILXISYNC” prefixes as described in this specification.

## Implementing IviLxiSync for an IVI-COM Instrument Driver

Guidelines for supporting the IviLxiSync API for an IVI-COM Instrument Driver are described in section 12.1.4. COM Interface Accessibility.

# IviLxiSync Repeated Capabilities

This specification makes extensive use of IVI repeated capabilities to model collections of arm and trigger alarms and sources, as well as outbound events.  Each of the five repeated capabilities defined in Section 2.1, Repeated Capability Names, shall be implemented according to the standard IVI repeated capability standards laid out in IVI 3.3-Standard Cross-Class Capabilities.  That specification, along with IVI 3.4-API Style Guide, describes three techniques for implementing IVI repeated capabilities – the selector technique, the parameter technique, and the IVI-COM collection technique.  All of the repeated capabilities in the IviLxiSync API shall use the IVI-COM collection technique for IVI-COM interfaces and the parameter technique for IVI-C interfaces.

## Repeated Capability Names

The IviLxiSync Specification defines five repeated capabilities.  Refer to the sections of IVI-3.1, Driver Architecture Specification that deal with repeated capabilities.  The relevant sections are Section 2.7, Repeated Capabilities, Section 4.1.9, Repeated Capabilities, Section 4.2.5, Repeated Capabilities, Section 4.3.9, Repeated Capabilities, and Section 5.9, Repeated Capability Identifiers and Selectors.

- IviLxiSyncArmAlarm

- IviLxiSyncArmSource

- IviLxiSyncTriggerAlarm

- IviLxiSyncTriggerSource

- IviLxiSyncEvent

### IviLxiSyncArmSource

In the configuration store, the name for the LXI arm source repeated capability shall be “IviLxiSyncArmSource”.

Although there is no intrinsic tie between Arm sources and Arm alarms, there is some potential for customer confusion since Trigger sources and Trigger alarms share the same namespace.  Therefore, drivers should place Arm sources and Arm alarms in the same namespace for consistency.

### IviLxiSyncArmAlarm

In the configuration store, the name for the LXI arm alarm repeated capability shall be “IviLxiSyncArmAlarm”.

Although there is no intrinsic tie between Arm sources and Arm alarms, there is some potential for customer confusion since Trigger sources and Trigger alarms share the same namespace.  Therefore, drivers should place Arm sources and Arm alarms in the same namespace for consistency.

### IviLxiSyncTriggerAlarm

In the configuration store, the name for the LXI trigger alarm capability shall be “IviLxiSyncTriggerAlarm”.

Note that the IviLxiSyncTriggerAlarm repeated capability group shares its namespace with IviLxiSyncTriggerSource.  See section 2.1.4.

### IviLxiSyncTriggerSource

In the configuration store, the name for the LXI trigger source capability shall be “IviLxiSyncTriggerSource”.

Note that the IviLxiSyncTriggerSource repeated capability shares its namespace with IviLxiSyncTriggerAlarm.  Therefore, once an Alarm has been defined, it may be used as a trigger source using the Trigger.TriggerSource attribute.

Other than this specific semantic connection, the repeated capability groups of TriggerSource and TriggerAlarm are unreleated.

### IviLxiSyncEvent

In the configuration store, the name for the LXI event repeated capability shall be “IviLxiSyncEvent”.

### Reserved Repeated Capability Identifiers

The following repeated capability identifiers are reserved and must be supported by the specific driver according to the LXI Functional Class (Class A, Class B, or Class C) of the instrument.  The specific driver must register each of the reserved repeated capability identifiers as physical repeated capability names in the IVI Configuration Store as per IVI-3.1: Driver Architecture Specification.

There are two types of reserved repeated capability identifiers.  The first is used for the pre-defined LXI trigger identifiers.  The second provides a predefined alarm to simplify programming alarms in the common case when only one is used.

The reserved trigger repeated capability identifiers apply only to the following repeated capabilities:

- IviLxiSyncArmSource

- IviLxiSyncTriggerSource

- IviLxiSyncEvent

The reserved alarm repeated capability identifier is “ALARM0”.  It applies to the following repeated capabilities:

- IviLxiSyncTriggerAlarm

- IviLxiSyncArmAlarm

Note that the LXI specification reserves event identifiers that begin with the characters “LXI” for LXI use.  The strings “LXI0”, “LXI1”, … ,”LXI7” refer to the 8 LXI wired trigger bus triggers.  See LXI 1.1 rule 6.4.5.

| Table 2-2. IviLxiSync Reserved Trigger Repeated Capability Identifiers |  |  |
| --- | --- | --- |
| Repeated Capability Identifier | Required for LXI Functional Class |  |
| LXI0 | Class A |  |
| LXI1 | Class A |  |
| LXI2 | Class A |  |
| LXI3 | Class A |  |
| LXI4 | Class A |  |
| LXI5 | Class A |  |
| LXI6 | Class A |  |
| LXI7 | Class A |  |
| LAN0 | Class A, Class B |  |
| LAN1 | Class A, Class B |  |
| LAN2 | Class A, Class B |  |
| LAN3 | Class A, Class B |  |
| LAN4 | Class A, Class B |  |
| LAN5 | Class A, Class B |  |
| LAN6 | Class A, Class B |  |
| LAN7 | Class A, Class B |  |

### Custom Repeated Capability Identifiers

In addition to the reserved repeated capability names defined in Section 2.1, many LXI devices allow custom repeated capability names for arm sources, trigger sources, and events.  These are referred to as custom repeated capability identifiers.

Custom repeated capability identifiers are used to dynamically add custom arm sources, arm alarms, trigger sources, trigger alarms, and events.  Client programs use standard Add and Remove methods (defined later in this specification) to manage the custom sources and events.

In addition to the requirements for valid IVI repeated capability identifiers, a custom repeated capability identifier must meet the following requirement:

The names must contain no more than 16 characters.

Since custom repeated capability identifiers are dynamic, they are not registered in the IVI Configuration Store when the specific driver is installed.

### Repeated Capability Identifier Case Sensitivity

Per IVI-3.1: Driver Architecture Specification, repeated capability identifiers are case sensitive.

### Repeated Capability Implementation Requirements

Drivers are not required to implement extensible repeated capabilities.  If a driver does not implement one of the following custom capability groups, it may define a fixed set of repeated capabilities and identifiers:

IviLxiSyncCustomTriggerSource

IviLxiSyncCustomArmSource

IviLxiSyncCustomTriggerAlarm

IviLxiSyncCustomerTriggerArm

IviLxiSyncCustomEvent

Note that for LXI instrument drivers the fixed set of repeated capabilities for IviLxiSyncTriggerSource and IviLxiSyncTriggerArm will generally include the reserved repeated capability identifiers listed in section 2.1.6 as well as others that may be relevant for that instrument.

## IviLxiSync Group Names

The IviLxiSync specification capabilities are divided up into a base capabilities group and multiple extension capabilities groups. This section defines names for each capability group and the functions and attributes that must be implemented by drivers that support the various capability groups.

### IviLxiSync Group Names

The capability groups for IviLxiSync are defined in the following table. The Group Name is used to represent a particular capability group.  It is returned as one of the possible group names from the drivers group capabilities attribute.  See IVI-3.3: Standard Cross Class Capabilities Specification.

| Table 2-1. IviLxiSync Group Names |  |  |
| --- | --- | --- |
| Group Name | Attributes | Functions |
| IviLxiSyncBase | Arm Count Arm Delay Arm Source Count Arm Source Detection Arm Source Enabled Arm Source EventId Arm Source Filter Arm Source Item (IVI-COM and IVI.NET Only) Arm Source Name (IVI-COM and IVI.NET Only) Arm Source Or Enabled Trigger Count Trigger Source Trigger Source Count Trigger Source Delay Trigger Source Detection Trigger Source EventId Trigger Source Item (IVI-COM and IVI.NET Only) Trigger Source Name (IVI-COM and IVI.NET Only) Trigger Source Filter | Configure Arm Source Disable All Arm Sources Get Arm Source Name (IVI-C Only) Configure Trigger Alarm Configure Trigger Source Disable All Trigger Alarms Get Trigger Source Name (IVI-C Only) |
| IviLxiSyncCustomArmSource |  | Add Arm Source Remove Arm Source Remove All Custom Arm Sources |
| IviLxiSyncCustomTriggerSource |  | Add Trigger Source Remove Trigger Source Remove All Custom Trigger Sources |
| IviLxiSyncTriggerAlarm | Trigger Alarm Count Trigger Alarm Enabled Trigger Alarm Item (IVI-COM Only) Trigger Alarm Name (IVI-COM Only) Trigger Alarm Period Trigger Alarm Repeat Count Trigger Alarm Time Seconds (IVI-C and IVI-COM Only) Trigger Alarm Time (IVI.NET Only) Trigger Alarm Time Fraction (IVI-C and IVI-COM Only) | Configure Trigger Alarm Disable All Trigger Alarms Get Trigger Alarm Name (IVI-C Only) |
| IviLxiSyncArmAlarm | Arm Alarm Count Arm Alarm Enabled Arm Alarm Item (IVI-COM and IVI.NET Only) Arm Alarm Name (IVI-COM and IVI.NET Only) Arm Alarm Period Arm Alarm Repeat Count Arm Alarm Time Seconds (IVI-C and IVI-COM Only) Arm Alarm Time (IVI.NET Only) Arm Alarm Time Fraction (IVI-C and IVI-COM Only) | Configure Arm Alarm Disable All Arm Alarms Get Arm Alarm Name (IVI-C Only) |
| IviLxiSyncCustomTriggerAlarm |  | Add Trigger Alarm Remove Trigger Alarm Remove All Trigger Alarms |
| IviLxiSyncCustomArmAlarm |  | Add Arm Alarm Remove Arm Alarm Remove All Custom Arm Alarms |
| IviLxiSyncEvent | Event Count Event Destination Path Event Drive Mode Event Item (IVI-COM and IVI.NET Only) Event Name (IVI-COM and IVI.NET Only) Event Slope Event Source Event Wired OR Bias Mode | Configure Event Disable All Events Get Event Name (IVI-C Only) |
| IviLxiSyncCustomEvent |  | Add Event Remove Event Remove All Custom Events |
| IviLxiSyncEventLog | Event Log Entry Count Event Log Enabled | Clear Event Log Entries Get Next Event Log Entry |
| IviLxiSyncSyncTime | Is Time Master Is Time Synchronized SystemTime (IVI.NET Only) | Get System Time (IVI-C and IVI-COM Only) |

## Boolean Attribute and Parameter Values

This specification uses True and False as the values for Boolean attributes and parameters. The following table defines the identifiers that are used for True and False in the IVI.NET, IVI-COM, and IVI-C architectures.

| Boolean Value | IVI.NET Identifier | IVI-COM Identifier | IVI-C Identifier |
| --- | --- | --- | --- |
| True | true | VARIANT_TRUE | VI_TRUE |
| False | false | VARIANT_FALSE | VI_FALSE |

## .NET Namespace

The .NET namespace for the IviLxiSync class is Ivi.LxiSync.

# IviLxiSyncArm Subsystem

The IviLxiSyncArm subsystem provides a mechanism for using inbound events to control when triggers are accepted by an LXI device.  Arming can be controlled via the LXI trigger bus, LXI-defined LAN events, or custom user-defined events.  In addition, arming can be controlled via a clocking mechanism.  Specifically, the LXI device can arm at a specific 1588 time and repeat at user-specified intervals.

The Arm subsystem also provides a way to logically OR or AND events together.  In cases where this is needed, the user can set the trigger to immediate and thereby initiate operations based on an AND or OR of events.

## Behavior Model

The figures below shows the behavior model for the IviLxiSyncArm subsystem.

Figure 2 shows how the IviLxiSync capabilities fit with the conventional IVI Initiate/Abort pattern.  Notice that  the Arm Clear message in Figure 3 occurs when the Initiate occurs.  Similarly, the arm is the number of arms that the instrument accepts before requiring another initiate.

Figure 3 shows the logic to selectively enable any of the inputs as well as selections for edge or level sensitivity and positive or negative slope.  The Arm Source Enabled attribute controls whether specific inputs are enabled, while Arm Source Edge and Arm Source Slope control edge or level detection and positive or negative slope, respectively.

Additional logic allows the input signals to be either OR-summed or AND-summed, as shown below.  The disposition of input signal summing is controlled by the Arm Source Or Enabled attribute.

The delay block shown in the diagram is used to compensate for electrical delay differences between the arm signal path and the signal path through the device under test.  Some devices, such as narrow-band crystal filters, have significant delay that must be taken into account in order to make valid measurements.

Arm capabilities should be implemented if the underlying instrument supports arming.  If the driver and instrument do not implement Arm capabilities, the existence of the functions in the class should not impact an application that only uses trigger.

Drivers that implement Arm should be designed so that users that do not use Arm capabilities do not need to issue any specific invocations of the driver to make the Arm system not intrude on a triggered measurement.  That is, if the application does not have a need for Arm, the default behaviors of the Arm system should configure the instrument so that it works with no required Arm preconditioning or additional programming.

Event

Logic

TriggerCounter >= End

TriggerCounter = 0

Trigger:

WaitingForTrigger = true

Trigger Event Detector

Wait: trigger delay

WaitingForTrigger = false

TriggerCounter++

Trigger Logic

N

Y

Measure:

Measuring = true   |   Acquire data   |   Measuring = false

ArmCounter >= End

ArmCounter = 0

ArmClear

Arm:

WaitingForArm = true

Arm Event Detector

Wait: arm delay

ArmCounter++

Arm Logic

N

Y

WaitingForArm = false

ArmClear

Init:Immediate

Init:Cont == On

Init:Cont == On

Idle:

Initiated:

Sweeping = true

N

Y

Y

OperationComplete = false

OperationComplete = true

Sweeping = false

Figure 2.  Overall Trigger Behavior Model

Figure 3.  IviLxiSyncArm Behavior Model

## IviLxiSyncArm Attributes

The IviLxiSyncArm subsystem defines the following attributes:

- Arm Count

- Arm Alarm Count

- Arm Alarm Enabled

- Arm Alarm Item (IVI-COM and IVI.NET Only)

- Arm Alarm Name (IVI-COM and IVI.NET Only)

- Arm Alarm Period

- Arm Alarm Repeat Count

- Arm Alarm Time (IVI.NET Only)

- Arm Alarm Time Fraction (IVI-C and IVI-COM Only)

- Arm Alarm Time Seconds (IVI-C and IVI-COM Only)

- Arm Delay

- Arm Source Count

- Arm Source Detection

- Arm Source Enabled

- Arm Source Event ID

- Arm Source Filter

- Arm Source Item (IVI-COM and IVI.NET Only)

- Arm Source Name (IVI-COM and IVI.NET Only)

- Arm Source Or Enabled

This section describes the behavior and requirements of each attribute. The actual value for each attribute ID is defined in Section 8, Attribute ID Definitions.

    - Arm Count

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | R/W | N/A | None | N/A |

.NET Property Name

Arm.ArmCount

COM Property Name

Arm.ArmCount

C Constant Name

IVILXISYNC_ATTR_ARM_COUNT

Description

Specifies the number of times the arm has to occur to complete the arm loop; that is, the number of arms that are accepted before the measurement must be initiated again.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Alarm Count

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | RO | IviLxiSyncArmAlarm | None | N/A |

.NET Property Name

Arm.Alarms.Count

This property is inherited from IIviRepeatedCapabilityCollection.

COM Property Name

Arm.Alarms.Count

C Constant Name

IVILXISYNC_ATTR_ARM_ALARM_COUNT

Description

Returns the number of currently available arm alarms.  The count returned includes the reserved repeated capability named “ALARM0” as defined Section 2.1.6, Reserved Repeated Capability Identifiers as well as any custom repeated capability identifiers.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Alarm Enabled

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViBoolean | R/W | IviLxiSyncArmAlarm | None | Configure Arm Alarm |

.NET Property Name

Arm.Alarms[].Enabled

COM Property Name

Arm.Alarms.Item().Enabled

C Constant Name

IVILXISYNC_ATTR_ARM_ALARM_ENABLED

Description

If set to True, the LXI device enables the arm alarm.  If set to False, the LXI device disables the arm alarm.

Compliance Notes

The specific driver shall implement both the True and False values.

Return Values

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Alarm Time Invalid | Error: Alarm time is inconsistent with the instrument hardware. For instance, the specified time (as indicated by the Arm Alarm Time Seconds and Arm Alarm Time Fractional attributes) has already passed. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Alarm Item (IVI-COM and IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| IIviLxiSyncArmAlarm* | RO | IviLxiSyncArmAlarm | None | N/A |

.NET Property Name

Arm.Alarms[String name]

This indexer is inherited from the base interface IIviRepeatedCapabilityCollection.  The name parameter uniquely identifies a particular arm alarm in the arm alarms collection.

COM Property Name

Arm.Alarms.Item ([in] BSTR SourceName)

C Constant Name

N/A

Description

Arm Alarm Item uniquely identifies an arm alarm in the arm alarms collection. It returns an interface pointer which can be used to control the attributes and other functionality of that arm alarm.

The Item property takes an arm alarm name.  If the user passes an invalid value for the source name parameter, the property returns an error.

Valid names include physical repeated capability identifiers and virtual repeated capability identifiers.

Return Values

If the IVI-COM driver cannot recognize the Name parameter, it returns an Unknown Name in Selector error as described in IVI-3.2: Inherent Capabilities Specification, Section 9.3.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Alarm Name (IVI-COM and IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | RO | IviLxiSyncArmAlarm | None | N/A |

.NET Property Name

Arm.Alarms[].Name

This property is inherited from IIviRepeatedCapabilityIdentification.

COM Property Name

Arm.Alarms.Name ([in] LONG SourceIndex)

C Constant Name

N/A.

(Use the GetArmAlarmName function.)

Description

Returns the physical repeated capability identifier defined by the specific driver for the arm alarm that corresponds to the one-based index that the user specifies. For custom arm sources added with the Add Arm Source function, this function returns the arm source name in the original casing used when Add Arm Source was called.

For C and COM, valid values for the SourceIndex parameter are between one and the value of the Arm Alarm Count attribute, inclusive. If the user passes an invalid value for the SourceIndex parameter, the value of this attribute is an empty string.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Alarm Period

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 (C/COM) | R/W | IviLxiSyncArmAlarm | None | Configure Arm Alarm |
| Ivi.Driver.PrecisionTimeSpan (.NET) | R/W | IviLxiSyncArmAlarm | None | Configure Arm Alarm |

.NET Property Name

Arm.Alarms[].Period

COM Property Name

Arm.Alarms.Item().Period

C Constant Name

IVILXISYNC_ATTR_ARM_ALARM_PERIOD

Description

Specifies the period of the arm alarm; that is, the amount of time that transpires before the alarm repeats.

For C and COM, time is in seconds, and a period of zero means there is no repeat and a single arm is generated.

For .NET, the units are implicit in the definition of PrecisionTimeSpan.  A period of PrecisionTimeSpan.Zero means there is no repeat and a single arm is generated.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Alarm Repeat Count

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | R/W | IviLxiSyncArmAlarm | None | Configure Arm Alarm |

.NET Property Name

Arm.Alarms[].RepeatCount

COM Property Name

Arm.Alarms.Item().RepeatCount

C Constant Name

IVILXISYNC_ATTR_ARM_ALARM_REPEAT_COUNT

Description

Specifies the number of times the trigger will occur at the period specified by the Arm Alarm Period attribute.  If Arm Alarm Repeat Count is zero, then the alarm shall repeat forever at the Arm Alarm Period.

For IVI-C, the defined value of IVILXISYNC_VAL_REPEAT_CONTINUOUS is provided to set the repeat count to forever.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Alarm Time (IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| Ivi.Driver.PrecisionDateTime | R/W | IviLxiSyncArmAlarm | None | Configure Arm Alarm |

.NET Property Name

Arm.Alarms[].Time

COM Property Name

N/A

(Use the Arm Alarm Time Seconds and Arm Alarm Time Fraction attributes.)

C Constant Name

N/A

(Use the Arm Alarm Time Seconds and Arm Alarm Time Fraction attributes.)

Description

Specifies the time at which the alarm will go off.  Note that PrecisionDateTime includes both seconds and fractional seconds, so that Arm Alarm Time in the .NET API replaces both Arm Alarm Time Seconds and Arm Alarm Time Fraction in the C/COM API.

Once the alarm goes off, it will repeat at the period set by Arm Alarm Period the number of times set by Arm Alarm Count.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Alarm Time Seconds (IVI-C and IVI-COM Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 | R/W | IviLxiSyncArmAlarm | None | Configure Arm Alarm |

.NET Property Name

N/A

(Use the Arm Alarm Time attribute.)

COM Property Name

Arm.Alarms.Item().TimeSeconds

C Constant Name

IVILXISYNC_ATTR_ARM_ALARM_TIME_SECONDS

Description

Specifies the seconds portion of time at which the alarm will go off.  Note that the actual time of the alarm is the sum of Arm Alarm Time Seconds and Arm Alarm Time Fraction.  The time is specified as the sum of two values because a single double-precision floating-point number does not have sufficient range and resolution to specify the time.

Once the alarm goes off, it will repeat at the period set by Arm Alarm Period the number of times set by Arm Alarm Count.

    - Arm Alarm Time Fraction (IVI-C and IVI-COM Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 | R/W | IviLxiSyncArmAlarm | None | Configure Arm Alarm |

.NET Property Name

N/A

(Use the Arm Alarm Time attribute.)

COM Property Name

Arm.Alarms.Item().TimeFraction

C Constant Name

IVILXISYNC_ATTR_ARM_ALARM_TIME_FRACTION

Description

Specifies the fractional portion of the time at which the alarm will go off.  Note that the actual time of the alarm is the sum of Arm Alarm Time Second and Arm Alarm Time Fraction.  The time is specified as the sum of two values because a single double-precision floating-point value does not have sufficient range and resolution to specify the time.

Once the alarm goes off, it will repeat at the period set by Arm Alarm Period the number of times set by Arm Alarm Count.

    - Arm Delay

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 (C/COM) | R/W | N/A | None | N/A |
| Ivi.Driver.PrecisionTimeSpan (.NET) | R/W | N/A | None | N/A |

.NET Property Name

Arm.Delay

COM Property Name

Arm.Delay

C Constant Name

IVILXISYNC_ATTR_ARM_DELAY

Description

Specifies the delay from when the arm logic satisfied until the waiting for trigger state is entered.  For C and COM, the units are seconds.  For .NET, the units are implicit in the definition of PrecisionTimeSpan.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Source Count

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | RO | IviLxiSyncArmSource | None | N/A |

.NET Property Name

Arm.Sources.Count

This property is inherited from IIviRepeatedCapabilityCollection.

COM Property Name

Arm.Sources.Count

C Constant Name

IVILXISYNC_ATTR_ARM_SOURCE_COUNT

Description

Returns the number of currently available arm sources.  The count returned includes any of the supported reserved repeated capability names defined in Section 2.1.6, Reserved Repeated Capability Identifiers as well as any custom repeated capability identifiers.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Source Detection

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | R/W | IviLxiSyncArmSource | None | Configure Arm Source |

.NET Property Name

Arm.Sources[].Detection

.NET Enumeration Name

Ivi.LxiSynchronization.ArmSourceDetection

COM Property Name

Arm.Sources.Item().Detection

COM Enumeration Name

IviLxiSyncSourceDetectionEnum

C Constant Name

IVILXISYNC_ATTR_ARM_SOURCE_DETECTION

Description

Specifies the style of arm source detection.

If the source is a LAN event and the source detection is set to rise, this Arm repeated capability will be satisfied when the designated LAN packet arrives with a True indication.  If the source detection is set to fall, this Arm repeated capability will be satisfied when a LAN packet arrives with a False indication.  If the detection is set to high, the source will be satisfied when the designated LAN packet arrives with a True indication and remain satisfied until the designated LAN packet arrives with a False indication.   If the detection is to low, the source will be satisfied when the designated LAN packet arrives with a False indication and remain satisfied until the designated LAN packet arrives with a True indication.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

Defined Values

| Compliance Notes Name | Description |  |  |
| --- | --- | --- | --- |
|  |  | Language | Identifier |
| Rise | Configures the LXI device to arm on the rising edge of the arm source. |  |  |
|  |  | .NET | ArmSourceDetection.Rise |
|  |  | C | IVILXISYNC_VAL_DETECTION_RISE |
|  |  | COM | IviLxiSyncArmSourceDetectionRise |
| Fall | Configures the LXI device to arm on the falling edge of the arm source. |  |  |
|  |  | .NET | ArmSourceDetection.Fall |
|  |  | C | IVILXISYNC_VAL_DETECTION_FALL |
|  |  | COM | IviLxiSyncArmSourceDetectionFall |
| High | Configures the LXI device to arm while the arm source is high, that is, while it remains true |  |  |
|  |  | .NET | ArmSourceDetection.High |
|  |  | C | IVILXISYNC_VAL_DETECTION_HIGH |
|  |  | COM | IviLxiSyncArmSourceDetectionHigh |
| Low | Configures the LXI device to arm while the arm source is low, that is, while it remains false |  |  |
|  |  | .NET | ArmSourceDetection.Low |
|  |  | C | IVILXISYNC_VAL_DETECTION_LOW |
|  |  | COM | IviLxiSyncArmSourceDetectionLow |

The specific driver shall implement all values.

    - Arm Source Enabled

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViBoolean | R/W | IviLxiSyncArmSource | None | Configure Arm Source |

.NET Property Name

Arm.Sources[].Enabled

COM Property Name

Arm.Sources.Item().Enabled

C Constant Name

IVILXISYNC_ATTR_ARM_SOURCE_ENABLED

Description

If set to True, the LXI device enables the arm source.. If set to False, the LXI device disables the arm source.  If a source is disabled, it has no affect on the summary arm signal.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

Compliance Notes

The specific driver shall implement both the True and False values.

    - Arm Source EventId

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | R/W | IviLxiSyncArmSource | None | N/A |

.NET Property Name

Arm.Sources[].EventId

COM Property Name

Arm.Sources.Item().EventId

C Constant Name

IVILXISYNC_ATTR_ARM_SOURCE_EVENTID

Description

This specifies the LAN event identifier that is associated with this arm source.  LAN Events with this identifier are accepted from  the source described in the filter.

The default value for EventId is the repeated capability specifier for this arm source.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Source Filter

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | R/W | IviLxiSyncArmSource | None | N/A |

.NET Property Name

Arm.Sources[].Filter

COM Property Name

Arm.Sources.Item().Filter

C Constant Name

IVILXISYNC_ATTR_ARM_SOURCE_FILTER

Description

Specifies a filter for restricting arm sources.  The filter specified by this attribute denotes the accepted sources. The syntax for specifying a filter is as follows:

<Filter> == [( <tcp> | <udp> | <any>) [, <Filter> ]]

<tcp> == <host> [:<port>]

<udp> == ALL  [: <port>]

<any> == : <port>

<host> is either a hostname or host number.  Note that the hostname can not be “ALL” since that would indicate the <udp> construct.

<port> is a series of decimal digits indicating the port number.

Specifying an empty string or VI_NULL means that LXI arm packets are accepted via either TCP or UDP multicast from any host.  Note that “:5044” is equivalent to the empty string since 5044 is the IANA registered port for LXI events (lxi-evntsvc).

In the <tcp>, <udp> and <any> constructs, <port> refers to the port the device receives the LAN message on.  If <port> is omitted from <tcp> or <udp>, packets are only accepted on the IANA registered port for LXI events (lxi-evntsvc).

Specifying the <host> (<tcp> construct) indicates that packets via TCP on the port indicated are accepted.

Specifying ALL (<udp> construct) indicates that UDP multicast packets are accepted if they are directed to the IANA registered port for LXI events (lxi-evntsvc)  on the IANA registered multicast address (LXI-EVENT).  No TCP packets are accepted unless a <tcp> syntax is also included in the filter.  The multicast address can not be altered with this syntax.

Specifying any protocol (<any> construct) indicates that both packets via TCP and UDP multicast packets are accepted if they are directed to the specified port.  UDP multicast packets must be received at the IANA registered multicast addres (LXI-EVENT).

The send port is not monitored.  This allows the transmitter to use any available port.

Drivers (and the corresponding instruments) that support this syntax are permitted to not support all possible filters syntaxes.

White space shall be ignored.  The <Filter> string is case insensitive.

Conventional devices should consider restricting the <port> to only the IANA registered port for LXI events (lxi-evntsvc) and not accepting the generalized syntax.

| Example Filter | Description |
| --- | --- |
| “192.168.0.1:23” | Accepts arm source packets via TCP on port 23 from the host at IP address 192.168.0.1. |
| “A_SIGGEN1:23,A_SPECAN2:23” | Accepts arm source packets via TCP on port 23 from the host with DNS name A_SIGGEN1 as well as packets via TCP on port 23 on the host with DNS name A_SPECAN2. |
| “192.168.0.1” | Accepts arm source packets from via TCP on the IANA registered port for LXI events (lxi-eventsvc) from the host at IP address 192.168.0.1. |
| “All:23,A_SPECAN2” | Accepts arm source packets via UDP multicast on the IANA assigned multicast address to port 23 from any host as well as packets via TCP on the IANA registered port for LXI events (lxi-eventsvc) from the host with DNS name A_SPECAN2. |
| “” | Accepts arm source packets via TCP or UDP multicast on the IANA assigned multicast address on the IANA registered port for LXI events (lxi-eventsvc) from any host. This is equivalent to <any> with the IANA registered port for LXI events (lxi-evntsvc). |
| “All” | Accepts arm source packets via UDP multicast on the IANA assigned multicast address on the IANA registered port for LXI events (lxi-eventsvc) from any host. This is equivalent to “”. |
| “All:8543” | Accepts arm source packets arriving via UDP multicast on the IANA UDP multicast address on port 8543 from any host. |
| “All, 192.168.1.1” | Accepts arm source packets arriving via UDP multicast on the IANA assigned multicast address at the IANA registered port for LXI events (lxi-eventsvc), or TCP packets arriving on the IANA registered port for LXI events (lxi-eventsvc) from IP address 192.168.1.1 |
| “:23” | Accepts arm source packets arriving via either UDP multicast or TCP on port 23 from any host. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Source Item (IVI-COM and IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| IIviLxiSyncArmSource* | RO | IviLxiSyncArmSource | None | N/A |

.NET Property Name

Arm.Sources[String name]

This indexer is inherited from IIviRepeatedCapabilityCollection.  The name parameter uniquely identifies a particular arm source in the arm sources collection.

COM Property Name

Arm.Sources.Item ([in] BSTR SourceName)

C Constant Name

N/A

Description

Arm Source Item uniquely identifies an arm source in the arm sources collection. It returns an interface pointer which can be used to control the attributes and other functionality of that arm source.

The Item property takes an arm source name. If the user passes an invalid value for the source name parameter, the property returns an error.

The .NET indexer may take an arm source name or an index into the arm source collection.  For .NET, valid values for the indexer’s index parameter are between zero and the value of the Arm Source Count attribute, minus 1, inclusive. If the user passes an invalid value for the name or index parameter, the indexer returns an error.

Valid names include physical repeated capability identifiers and virtual repeated capability identifiers.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Source Name (IVI-COM and IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | RO | IviLxiSyncArmSource | None | N/A |

.NET Property Name

Arm.Sources[].Name

This property is inherited from IIviRepeatedCapabilityIdentification.

COM Property Name

Arm.Sources.Name ([in] LONG SourceIndex)

C Constant Name

N/A.

(Use the Get Arm Source Name function.)

Description

Returns the physical repeated capability identifier defined by the specific driver for the arm source that corresponds to the index that the user specifies.  For custom arm sources added with the Add Arm Source function, this function returns the arm source name in the original casing used when Add Arm Source was called.

For C and COM, valid values for the SourceIndex parameter are between one and the value of the Arm Source Count attribute, inclusive. If the user passes an invalid value for the SourceIndex parameter, the value of this attribute is an empty string.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Arm Source Or Enabled

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViBoolean | R/W | N/A | None | N/A |

.NET Property Name

Arm.Sources.OrEnabled

COM Property Name

Arm.Sources.OrEnabled

C Constant Name

IVILXISYNC_ATTR_ARM_SOURCE_OR_ENABLED

Description

If set to True, the arm sources use OR-summing of the arm sources.  If set to False, the arm sources use AND-summing of the arm sources.

Compliance Notes

The specific driver shall implement both the True and False values.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

## IviLxiSyncArm Functions

The IviLxiSyncArm subsystem defines the following functions:

- Add Arm Alarm

- Add Arm Source

- Configure Arm Alarm

- Configure Arm Source

- Disable All Arm Alarms

- Disable All Arm Sources

- Get Arm Alarm Name (IVI-C Only)

- Get Arm Source Name (IVI-C Only)

- Remove Arm Alarm

- Remove Arm Source

- Remove All Custom Arm Alarms

- Remove All Custom Arm Sources

This section describes the behavior and requirements of each function.

    - Add Arm Alarm

Description

This function creates a new arm alarm.

When a new alarm is added, the default values for the associated attributes are as follows:

| Attribute | Default Value |
| --- | --- |
| Arm Alarm Enabled | False |
| Arm Alarm Period | 0 |
| Arm Alarm Repeat Count | 1 |
| Arm Alarm Time (.NET) | 00:00:00 Jan 1, 1970 |
| Arm Alarm Time Seconds (C/COM) | 0 |
| Arm Alarm Time Fraction (C/COM) | 0 |

.NET Method Prototype

IIviLxiSyncArmAlarm Arm.Alarms.Add(String alarmName);

COM Method Prototype

HRESULT Arm.Alarms.Add([in] BSTR AlarmName);

C Prototype

ViStatus IviLxiSync_AddArmAlarm (ViSession Vi, 
                              ViConstString AlarmName);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| AlarmName | Specifies the name of the arm alarm to create. | ViConstString |
| Outputs | Description | Base Type |
| Return value (.NET) | A reference to the arm alarm object that was added . | IIviLxiSyncArmAlarm |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Alarm Exists | Error: Alarm already exists. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

The table below specifies additional class-defined exceptions for this method.

| Exception Class | Description |
| --- | --- |
| AlarmExistsException | Alarm already exists. |

    - Add Arm Source

Description

This function creates a new arm source.

When a new source is added, the default values for the associated attributes are as follows:

| Attribute | Default Value |
| --- | --- |
| Arm Source Edge | True (edge-detection) |
| Arm Source Enabled | True |
| Arm Source EventId | The repeated capability source name specified in this call |
| Arm Source Slope | Positive |
| Arm Source Filter | “” (Empty string) |
| Arm Source Or Enabled | False |

.NET Method Prototype

IIviLxiSyncArmSource Arm.Sources.Add(String sourceName);

COM Method Prototype

HRESULT Arm.Sources.Add([in] BSTR SourceName);

C Prototype

ViStatus IviLxiSync_AddArmSource (ViSession Vi, 
                               ViConstString SourceName);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| SourceName | Specifies the name of the arm source to create. | ViConstString |
| Outputs | Description | Base Type |
| Return value (.NET) | A reference to the arm source object that was added . | IIviLxiSyncArmAlarm |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Event Source Exists | Error: Event source already exists. |
| Out of Event Resources | Error: Out of event resources. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

The table below specifies additional class-defined exceptions for this method.

| Exception Class | Description |
| --- | --- |
| EventSourceExistsException | Event source already exists. |
| OutOfEventResourcesException | Out of event resources. |

    - Configure Arm Alarm

Description

This function configures the most commonly used attributes of the arm alarm subsystem.

.NET Method Prototype

void Arm.Alarms[].Configure(Boolean enabled,
                           PrecisionDateTime time,
                           PrecisionTimeSpan period, 
                           Int32 repeatCount);

COM Method Prototype

HRESULT Arm.Alarms.Item().Configure([in] VARIANT_BOOL Enabled,
                                    [in] DOUBLE TimeSeconds,
                                    [in] DOUBLE TimeFraction, 
                                    [in] DOUBLE Period, 
                                    [in] LONG RepeatCount);

C Prototype

ViStatus IviLxiSync_ConfigureArmAlarm (ViSession Vi, 
                                       ViConstString AlarmName, 
                                       ViBoolean Enabled,
                                       ViReal64 TimeSeconds, 
                                       ViReal64 TimeFraction, 
                                       ViReal64 Period, 
                                       ViInt32 RepeatCount);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| AlarmName | The name of the alarm. | ViConstString |
| Enabled | Enables or disables the arm alarm. The driver uses this value to set the Arm Alarm Enabled attribute. See the attribute description for more information. | ViBoolean |
| time (.NET) | PrecisionDateTime. The driver uses this value to set the Arm Alarm Time attribute. See the attribute description for more information. | Ivi.Driver. PrecisionDateTime |
| TimeSeconds (C/COM) | Specifies the seconds part of the time. The driver uses this value to set the Arm Alarm Time Seconds attribute. See the attribute description for more information. | ViReal64 |
| TimeFraction (C/COM) | Specifies the fractional part of the time. The driver uses this value to set the Arm Alarm Time Fractional attribute. See the attribute description for more information. | ViReal64 |
| Period | Specifies the period of the arm alarm. The driver uses this value to set the Arm Alarm Period attribute. See the attribute description for more information. | ViReal64 (C/COM) Ivi.Driver. PrecisionTimeSpan (.NET) |
| RepeatCount | Specifies the number of times to repeat the trigger at the period specified by the Arm Alarm Repeat Period attribute. The driver uses this value to set the Arm Alarm Repeat Count attribute. See the attribute description for more information. | ViInt32 |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Configure Arm Source

Description

This function configures the most commonly used attributes of the arm source subsystem.

.NET Method Prototype

void Arm.Sources[].Configure(Boolean enabled,
                            ArmSourceDetection detection);

COM Method Prototype

HRESULT Arm.Sources.Item().Configure([in] VARIANT_BOOL Enabled,
                                     [in] IviLxiSyncArmSourceDetectionEnum Detection);

C Prototype

ViStatus IviLxiSync_ConfigureArmSource (ViSession Vi, 
                                        ViConstString SourceName, 
                                        ViBoolean Enabled, 
                                        ViInt32 Detection);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| SourceName | Name of the arm source to configure. | ViConstString |
| Enabled | Enables or disables the arm source. The driver uses this value to set the Arm Source Enabled attribute. See the attribute description for more information. | ViBoolean |
| Detection | Specifies the detection mode of the arm source. The driver uses this value to set the Arm Source Detection attribute. See the attribute description for more information. | ViInt32 |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Disable All Arm Alarms

Description

This function disables all arm alarms.  The specific driver uses this function to set the Arm Alarm Enabled property to False for all arm alarms.

.NET Method Prototype

void Arm.Alarms.DisableAll();

COM Method Prototype

HRESULT Arm.Alarms.DisableAll();

C Prototype

ViStatus IviLxiSync_DisableAllArmAlarms (ViSession Vi);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Disable All Arm Sources

Description

This function disables all arm sources.  The specific driver uses this function to set the Arm Source Enabled property to False for all arm sources.

.NET Method Prototype

void Arm.Sources.DisableAll();

COM Method Prototype

HRESULT Arm.Sources.DisableAll();

C Prototype

ViStatus IviLxiSync_DisableAllArmSources (ViSession Vi);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return..NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Get Arm Alarm Name (IVI-C Only)

Description

This function returns the physical repeated capability identifier that corresponds to the one-based index that the user specifies. If the value that the user passes for the AlarmIndex parameter is less than one or greater than the value of the Arm Alarm Count attribute, the function returns an empty string in the AlarmName parameter and returns an error.  For custom arm sources added vith the Add Arm Source function, this function returns the arm source name in the original casing used when Add Arm Source was called.

.NET Method Prototype

N/A

(Use the Arm.Alarms[].Name property)

COM Method Prototype

N/A

(Use the Arm.Alarms.Item().Name property.)

C Prototype

ViStatus IviLxiSync_GetArmAlarmName (ViSession Vi, 
                                     ViInt32 AlarmIndex, 
                                     ViInt32 AlarmNameBufferSize, 
                                     ViChar AlarmName[]);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| AlarmIndex | A one-based index that defines which name to return. | ViInt32 |
| AlarmNameBufferSize | The number of bytes in the ViChar array that the user specifies for the AlarmName parameter. | ViInt32 |
| Outputs | Description | Data Type |
| AlarmName | The buffer into which the function returns the alarm name that corresponds to the index the user specifies. The caller may pass VI_NULL for this parameter if the AlarmNameBufferSize parameter is 0. | ViChar[] |

Return Values (C)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

    - Get Arm Source Name (IVI-C Only)

Description

This function returns the physical repeated capability identifier that corresponds to the one-based index that the user specifies. If the value that the user passes for the SourceIndex parameter is less than one or greater than the value of the Arm Source Count attribute, the function returns an empty string in the SourceName parameter and returns an error.  For custom arm sources added vith the Add Arm Source function, this function returns the arm source name in the original casing used when Add Arm Source was called.

.NET Method Prototype

N/A

(Use the Arm.Sources[].Name property)

COM Method Prototype

N/A

(Use the Arm.Sources.Item().Name property.)

C Prototype

ViStatus IviLxiSync_GetArmSourceName (ViSession Vi, 
                                      ViInt32 SourceIndex, 
                                      ViInt32 SourceNameBufferSize, 
                                      ViChar SourceName[]);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| SourceIndex | A one-based index that defines which name to return. | ViInt32 |
| SourceNameBufferSize | The number of bytes in the ViChar array that the user specifies for the SourceName parameter. | ViInt32 |
| Outputs | Description | Data Type |
| SourceName | The buffer into which the function returns the source name that corresponds to the index the user specifies. The caller may pass VI_NULL for this parameter if the SourceNameBufferSize parameter is 0. | ViChar[] |

Return Values (C)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

    - Remove Arm Alarm

Description

This function removes an arm alarm.  Any resources associated with this alarm should be freed.

The AlarmName parameter is case-insensitive.

.NET Method Prototype

void Arm.Alarms.Remove(String alarmName);

COM Method Prototype

HRESULT Arm.Alarms.Remove([in] BSTR AlarmName);

C Prototype

ViStatus IviLxiSync_RemoveArmAlarm (ViSession Vi, 
                                 ViConstString AlarmName);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| AlarmName | Specifies the name of the arm alarm to remove. | ViConstString |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Alarm Does Not Exist | Error: Alarm does not exist. |
| Attempt To Remove Reserved Repeated Capability | Error: The repeated capability name is reserved and cannot be removed. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

The table below specifies additional class-defined exceptions for this method.

| Exception Class | Description |
| --- | --- |
| AlarmDoesNotExistException | Alarm does not exist. |
| CannotRemoveReservedRepeatedCapabilityException | The repeated capability name is reserved and cannot be removed. |

    - Remove Arm Source

Description

This function removes an arm source.  Any resources associated with this alarm should be freed.

The SourceName parameter is case-insensitive.

.NET Method Prototype

void Arm.Sources.Remove(String sourceName);

COM Method Prototype

HRESULT Arm.Sources.Remove([in] BSTR SourceName);

C Prototype

ViStatus IviLxiSync_RemoveArmSource (ViSession Vi, 
                                     ViConstString SourceName);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| SourceName | Specifies the name of the arm source to remove. | ViConstString |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Event Source Does Not Exist | Error: Event source does not exist. |
| Attempt To Remove Reserved Repeated Capability | Error: The repeated capability name is reserved and cannot be removed. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

The table below specifies additional class-defined exceptions for this method.

| Exception Class | Description |
| --- | --- |
| EventSourceDoesNotExistException | Event source does not exist. |
| CannotRemoveReservedRepeatedCapabilityException | The repeated capability name is reserved and cannot be removed. |

    - Remove All Custom Arm Alarms

Description

This function removes all of the arm alarms that were added using the Add Arm Alarm function.

.NET Method Prototype

void Arm.Alarms.RemoveAllCustomArmAlarms();

COM Method Prototype

HRESULT Arm.Alarms.RemoveAllCustomArmAlarms();

C Prototype

ViStatus IviLxiSync_RemoveAllCustomArmAlarms (ViSession Vi);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Remove All Custom Arm Sources

Description

This function removes all of the custom arm sources that were added using the Add Arm Source function.  The arm sources associated with the reserved repeated capability identifiers, as defined in Section 2.1.6, Reserved Repeated Capability Identifiers, are not affected by this function.

.NET Method Prototype

void Arm.Sources.RemoveAllCustomArmSources();

COM Method Prototype

HRESULT Arm.Sources.RemoveAllCustomArmSources();

C Prototype

ViStatus IviLxiSync_RemoveAllCustomArmSources (ViSession Vi);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

# IviLxiSyncTrigger Subsystem

The IIviLxiSyncTrigger interface provides a mechanism for using inbound events to control when to trigger an LXI device.  This subsystem provides control for triggers from the LXI trigger bus, LXI-defined LAN events, or custom user-defined events.  Methods and properties for controlling this mode of triggering are exposed via the IIviLxiSyncTriggerSources sub-interface.  In addition, triggering can be controlled via a clocking mechanism.  Specifically, the LXI device can trigger at a specific time or at user-specified intervals.

## Behavior Model

The IviLxiSyncTrigger subsystem behavior model is somewhat simpler than that of the IviLxiSyncArm subsystem.  This is due to the fact that only a single trigger source is selected at a time and that triggers are always initiated on an edge (that is, when the trigger occurs, the operation begins immediately, so a “level trigger” is not a meaningful concept).

If an application needs to OR-sum or AND-sum events and use the result to trigger a measurement (or other operation).  This can be done by using the summing feature of the IviLxiSyncArm subsystem and selecting an Immediate trigger in the trigger logic.  A similar technique can be used for trigger gating (controlling the acquisition process with the level of a signal).  To do this set the Arm Source Edge attribute to False and select the Immediate trigger in the trigger logic.

The Trigger subsystem should be designed so that users that do not use these trigger capabilities do not need to issue any specific invocations of the driver to prevent the trigger system from intruding on a measurement that is made with Initiate or Read.  That is, if the application does not have a need for Trigger, the default behaviors of the Trigger system should configure the instrument so that it works with no required Trigger conditions beyond the Initiate.

Figure 4.  IviLxiSyncTrigger Behavior Model

## IviLxiSyncTrigger Attributes

The IviLxiSyncTrigger subsystem defines the following attributes:

- Trigger Alarm Count

- Trigger Alarm Enabled

- Trigger Alarm Item (IVI-COM and IVI.NET Only)

- Trigger Alarm Name (IVI-COM and IVI.NET Only)

- Trigger Alarm Period

- Trigger Alarm Repeat Count

- Trigger Alarm Time  (IVI.NET Only)

- Trigger Alarm Time Fraction (IVI-C and IVI-COM Only)

- Trigger Alarm Time Seconds (IVI-C and IVI-COM Only)

- Trigger Count

- Trigger Source

- Trigger Source Count

- Trigger Source Delay

- Trigger Source Detection

- Trigger Source Event ID

- Trigger Source Item (IVI-COM and IVI.NET Only)

- Trigger Source Name (IVI-COM and IVI.NET Only)

- Trigger Source Filter

This section describes the behavior and requirements of each attribute. The actual value for each attribute ID is defined in Section 8, Attribute ID Definitions.

    - Trigger Alarm Count

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | RO | IviLxiSyncTriggerAlarm | None | N/A |

.NET Property Name

Trigger.Alarms.Count

COM Property Name

Trigger.Alarms.Count

C Constant Name

IVILXISYNC_ATTR_TRIGGER_ALARM_COUNT

Description

Returns the number of currently available trigger alarms.  The count returned includes the reserved repeated capability named “ALARM0” as defined Section 2.1.6, Reserved Repeated Capability Identifiers as well as any custom repeated capability identifiers.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Alarm Enabled

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViBoolean | R/W | IviLxiSyncTriggerAlarm | None | Configure Trigger Alarm |

.NET Property Name

Trigger.Alarms[].Enabled

COM Property Name

Trigger.Alarms.Item().Enabled

C Constant Name

IVILXISYNC_ATTR_TRIGGER_ALARM_ENABLED

Description

If set to True, the LXI device enables the trigger alarm.  If set to False, the LXI device disables the trigger alarm.

Compliance Notes

The specific driver shall implement both the True and False values.

Return Values

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Alarm Time Invalid | Error: Alarm time is inconsistent with the instrument hardware. For instance, the specified time (as indicated by the Trigger Alarm Time Seconds and Trigger Alarm Time Fractional attributes) has already passed. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Alarm Item (IVI-COM and IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| IIviLxiSyncTriggerAlarm* | RO | IviLxiSyncArmAlarm | None | N/A |

.NET Property Name

Trigger.Alarms[String name]

This indexer is inherited from IIviRepeatedCapabilityCollection.  The name parameter uniquely identifies a particular trigger alarm in the Trigger Alarms collection.

COM Property Name

Trigger.Alarms.Item ([in] BSTR SourceName)

C Constant Name

N/A

Description

Trigger Alarm Item uniquely identifies a trigger alarm in the trigger alarms collection. It returns an interface pointer which can be used to control the attributes and other functionality of that trigger alarm.

The Item property takes a trigger alarm name.  If the user passes an invalid value for the SourceName parameter, the property returns an error.

Valid names include physical repeated capability identifiers and virtual repeated capability identifiers.

Return Values

If the IVI-COM driver cannot recognize the Name parameter, it returns an Unknown Name in Selector completion code as described in IVI-3.2: Inherent Capabilities Specification, Section 9.3.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Alarm Name (IVI-COM and IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | RO | IviLxiSyncTriggerAlarm | None | N/A |

.NET Property Name

Trigger.Alarms[].Name

This property is inherited from IIviRepeatedCapabilityIdentification.

COM Property Name

Trigger.Alarms.Name ([in] LONG SourceIndex)

C Constant Name

N/A

(Use the GetTriggerAlarmName function.)

Description

Returns the physical repeated capability identifier defined by the specific driver for the trigger alarm that corresponds to the index that the user specifies. For custom arm sources added with the Add Trigger Source function, this function returns the arm source name in the original casing used when Add Trigger Source was called.

For C and COM, valid values for the SourceIndex parameter are between one and the value of the Trigger Alarm Count attribute, inclusive. If the user passes an invalid value for the SourceIndex parameter, the value of this attribute is an empty string.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Alarm Period

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 (C/COM) | R/W | N/A | None | Configure Trigger Alarm |
| Ivi.Driver.PrecisionTimeSpan (.NET) | R/W | N/A | None | Configure Trigger Alarm |

.NET Property Name

Trigger.Alarms[].Period

COM Property Name

Trigger.Alarms.Item().Period

C Constant Name

IVILXISYNC_ATTR_TRIGGER_ALARM_PERIOD

Description

Specifies the period of the trigger alarm; that is, the amount of time that transpires before the alarm repeats.

For C and COM, time is in seconds, and a period of zero means there is no repeat and a single trigger is generated.

For .NET, the units are implicit in the definition of PrecisionTimeSpan.  A period of PrecisionTimeSpan.Zero means there is no repeat and a single arm is generated.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Alarm Repeat Count

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | R/W | N/A | None | Configure Trigger Alarm |

.NET Property Name

Trigger.Alarms[].RepeatCount

COM Property Name

Trigger.Alarms.Item().RepeatCount

C Constant Name

IVILXISYNC_ATTR_TRIGGER_ALARM_REPEAT_COUNT

Description

Specifies the number of times the trigger will occur at the period specified by the Trigger Alarm Period attribute.  If Trigger Alarm Repeat Period is non-zero and Trigger Alarm Repeat Count is zero, then the alarm shall repeat forever at the Trigger Alarm Period.

For IVI-C, the defined value of IVILXISYNC_VAL_REPEAT_CONTINUOUS is provided to set the repeat count to forever.

    - Trigger Alarm Time (IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| PrecisionDateTime | R/W | N/A | None | Configure Trigger Alarm |

.NET Property Name

Trigger.Alarms[].Time

COM Property Name

N/A

(Use the Trigger Alarm Time Seconds and Trigger Alarm Time Fraction attributes.)

C Constant Name

N/A

(Use the Trigger Alarm Time Seconds and Trigger Alarm Time Fraction attributes.)

Description

Specifies the time at which the alarm will go off.  Note that PrecisionDateTime includes both seconds and fractional seconds, so that Arm Alarm Time in the .NET API replaces both Trigger Alarm Time Seconds and Trigger Alarm Time Fraction in the C/COM API.

Once the alarm goes off, it will repeat at the period set by Trigger Alarm Period the number of times set by Trigger Alarm Count.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Alarm Time Seconds (IVI-C and IVI-COM Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 | R/W | N/A | None | Configure Trigger Alarm |

.NET Property Name

N/A

(Use the Trigger Alarm Time attribute.)

COM Property Name

Trigger.Alarms.Item().TimeSeconds

C Constant Name

IVILXISYNC_ATTR_TRIGGER_ALARM_TIME_SECONDS

Description

Specifies the seconds portion of the time at which the alarm will go off.  Note that the actual time of the alarm is the sum of Trigger Alarm Time Seconds and Trigger Alarm Time Fraction.  The time is specified as the sum of two values because a single double-precision floating-point does not have sufficient range and resolution to specify the time.

Once the alarm goes off, it will repeat at the period set by Trigger Alarm Period the number of times set by Trigger Alarm Count.

    - Trigger Alarm Time Fraction (IVI-C and IVI-COM Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 | R/W | N/A | None | Configure Trigger Alarm |

.NET Property Name

N/A

(Use the Trigger Alarm Time attribute.)

COM Property Name

Trigger.Alarms.Item().TimeFraction

COM Enumeration Name

N/A

C Constant Name

IVILXISYNC_ATTR_TRIGGER_ALARM_TIME_FRACTION

Description

Specifies the fractional seconds portion of the time at which the alarm will go off.  Note that the actual time of the alarm is the sum of Trigger Alarm Time Seconds and Trigger Alarm Time Fraction.  The time is specified as the sum of two values because a single double-precision floating-point value does not have sufficient range and resolution to specify the time.

Once the alarm goes off, it will repeat at the period set by Trigger Alarm Period the number of times set by Trigger Alarm Count.

    - Trigger Count

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | R/W | N/A | None | N/A |

.NET Property Name

Trigger.TriggerCount

COM Property Name

Trigger.TriggerCount

C Constant Name

IVILXISYNC_ATTR_TRIGGER_COUNT

Description

Specifies the number of times a trigger has to occur to complete a measurement; that is, the number of triggers that are accepted before the measurement must be armed again.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Source

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | R/W | N/A | None | N/A |

.NET Property Name

Trigger.TriggerSource

COM Property Name

Trigger.TriggerSource

C Constant Name

IVILXISYNC_ATTR_TRIGGER_SOURCE

Description

Specifies which of the available trigger sources to use as the signal for triggering the device-specific operation (for example, a measurement).

The value specified for this attribute may be one of the names in the IviLxiSyncTriggerSource repeated capability collection as returned from either the Trigger Source Name attribute (for IVI-COM and IVI.NET) or the GetTriggerSourceName function (for IVI-C).

The value specified for this attribute may also be one of the names in the IviLxiSyncTriggerAlarm repeated capability collection as returned from either the Trigger Alarm Name attribute (for IVI-COM and IVI.NET) or the GetTriggerAlarmName function (for IVI-C).

The name specified here may also correspond to a non-LXI trigger event.  For instance, the caller can use this attribute to program the trigger source to external or immediate triggering, by specifying values such as “EXT” or “INT”.  Such trigger source names are device-dependent.

If the device trigger source has been programmed to a non-LXI event using an attribute or function other than the Trigger Source attribute, then this attribute shall return that value when read.  For instance, if the specific driver implements an IVI instrument class and the class-compliant API is used to set the trigger source to external, then this property shall return a string that reflects the value set through the class-compliant API.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Source Count

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | RO | IviLxiSyncTriggerSource | None | N/A |

.NET Property Name

Trigger.Sources.Count

This property is inherited from IIviRepeatedCapabilityCollection.

COM Property Name

Trigger.Sources.Count

C Constant Name

IVILXISYNC_ATTR_TRIGGER_SOURCE_COUNT

Description

Returns the number of currently available trigger sources.  The count returned includes any of the supported reserved repeated capability names defined in Section 2.1.6, Reserved Repeated Capability Identifiers as well as any custom repeated capability identifiers.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Source Delay

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 (C/COM) | R/W | IviLxiSyncTriggerSource | None | Configure Trigger Source |
| Ivi.Driver.PrecisionTimeSpan (.NET) | R/W | IviLxiSyncTriggerSource | None | Configure Trigger Source |

.NET Property Name

Trigger.Sources[].Delay

COM Property Name

Trigger.Sources.Item().Delay

C Constant Name

IVILXISYNC_ATTR_TRIGGER_SOURCE_DELAY

Description

Specifies the trigger source delay from when the trigger logic is satisfied until the device specific action (for instance a measurement) is triggered.  A negative value implies pre-trigger acquisition. For C and COM, the units are seconds.  For .NET, the units are implicit in the definition of PrecisionTimeSpan.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Source Detection

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | R/W | IviLxiSyncTriggerSource | None | Configure Trigger Source |

.NET Property Name

Trigger.Sources[].Detection

.NET Enumeration Name

Slope

COM Property Name

Trigger.Sources.Item().Detection

COM Enumeration Name

IviLxiSyncSourceDetectionEnum

C Constant Name

IVILXISYNC_ATTR_TRIGGER_SOURCE_DETECTION

Description

Specifies the slope of the trigger source.

If the source is a LAN event and the source slope is set to positive, this Trigger repeated capability will be satisfied when the designated LAN packet arrives with a true indication.  If the source slope is set to negative, this Trigger repeated capability will be satisfied when a LAN packet arrives with a false indication.

Defined Values

| .NET Exceptions Name | Description |  |  |
| --- | --- | --- | --- |
|  |  | Language | Identifier |
| Rise | Configures the LXI device to trigger on the rising edge of the trigger source. |  |  |
|  |  | .NET | Slope.Positive |
|  |  | C | IVILXISYNC_VAL_DETECTION_RISE |
|  |  | COM | IviLxiSyncTriggerSourceDetectionRise |
| Fall | Configures the LXI device to trigger on the falling edge of the trigger source. |  |  |
|  |  | .NET | Slope.Negative |
|  |  | C | IVILXISYNC_VAL_DETECTION_FALL |
|  |  | COM | IviLxiSyncTriggerSourceDetectionFall |

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Source EventId

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | R/W | IviLxiSyncArmSource | None | N/A |

.NET Property Name

Trigger.Sources[].EventId

COM Property Name

Trigger.Sources.Item().EventId

C Constant Name

IVILXISYNC_ATTR_ARM_SOURCE_EVENTID

Description

This specifies the LAN event identifier that is associated with this trigger source.  LAN Events with this identifier are accepted from  the source described in the filter.

The default value for EventId is the repeated capability specifier for this trigger source.

    - Trigger Source Item (IVI-COM and IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| IIviLxiSyncTriggerSource* | RO | IviLxiSyncArmSource | None | N/A |

.NET Property Name

Trigger.Sources[String name]

This indexer is inherited from IIviRepeatedCapabilityCollection.  The name parameter uniquely identifies a particular trigger source in the trigger sources collection.

COM Property Name

Trigger.Sources.Item ([in] BSTR SourceName)

C Constant Name

N/A

Description

Trigger Source Item uniquely identifies a trigger source in the trigger sources collection. It returns an interface pointer which can be used to control the attributes and other functionality of that trigger source.

The Item property takes a trigger source name.  If the user passes an invalid value for the SourceName parameter, the property returns an error.

Valid names include physical repeated capability identifiers and virtual repeated capability identifiers.

Return Values

If the IVI-COM driver cannot recognize the Name parameter, it returns an Unknown Name in Selector completion code as described in IVI-3.2: Inherent Capabilities Specification, Section 9.3.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Source Name (IVI-COM and IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | RO | IviLxiSyncTriggerSource | None | N/A |

.NET Property Name

Trigger.Sources[].Name

This property is inherited from IIviRepeatedCapabilityIdentification.

COM Property Name

Trigger.Sources.Name ([in] LONG SourceIndex)

C Constant Name

N/A

(Use the GetTriggerSourceName function.)

Description

Returns the physical repeated capability identifier defined by the specific driver for the trigger source that corresponds to the one-based index that the user specifies.  For custom trigger sources added vith the Add Trigger Source function, this function returns the trigger source name in the original casing used when Add Trigger Source was called.

For C and COM, valid values for the SourceIndex parameter are between one and the value of the Trigger Source Count attribute. If the user passes an invalid value for the SourceIndex parameter, the value of this attribute is an empty string.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Trigger Source Filter

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | R/W | IviLxiSyncTriggerSource | None | N/A |

.NET Property Name

Trigger.Sources[].Filter

COM Property Name

Trigger.Sources.Item().Filter

C Constant Name

IVILXISYNC_ATTR_TRIGGER_SOURCE_FILTER

Description

Specifies a filter for restricting trigger sources.  The filter specified by this attribute denotes the accepted sources. The syntax for specifying a filter is as follows:

<Filter> == [( <tcp> | <udp> | <any>) [, <Filter> ]]

<tcp> == <host> [:<port>]

<udp> == ALL  [: <port>]

<any> == : <port>

<host> is either a hostname or host number.  Note that the hostname can not be “ALL” since that would indicate the <udp> construct.

<port> is a series of decimal digits indicating the port number.

Specifying an empty string or VI_NULL means that LXI trigger packets are accepted via either TCP or UDP multicast from any host.  “:5044” is equivalent to the empty string since 5044 is the IANA registered port for LXI events (lxi-evntsvc).

In the <tcp>, <udp> and <any> constructs, <port> refers to the port the device receives the LAN message on.  If <port> is omitted from <tcp> or <udp>, packets are only accepted on the IANA registered port for LXI events (lxi-evntsvc).

Specifying the <host> (<tcp> construct) indicates that packets via TCP on the port indicated are accepted.

Specifying ALL (<udp> construct) indicates that UDP multicast packets are accepted if they are directed to the IANA registered port for LXI events (lxi-evntsvc)  on the IANA registered multicast address (LXI-EVENT).  No TCP packets are accepted unless a <tcp> syntax is also included in the filter.  The multicast address can not be altered with this syntax.

Specifying any protocol (<any> construct) indicates that both TCP and UDP multicast packets are accepted if they are directed to the specified port.  UDP multicast packets must be received at the IANA registered multicast addres (LXI-EVENT).

The send port is not monitored.  This allows the transmitter to use any available port.

Drivers (and the corresponding instruments) that support this syntax are permitted to not support all possible filters syntaxes.

White space shall be ignored.  The <Filter> string is case insensitive.

Conventional devices should consider restricting the <port> to only the IANA registered port for LXI events (lxi-eventsvc) and not accepting the generalized syntax.

| Example Filter | Description |
| --- | --- |
| “192.168.0.1:23” | Accepts trigger source packets via TCP on port 23 from the host at IP address 192.168.0.1. |
| “A_SIGGEN1:23,A_SPECAN2:23” | Accepts trigger source packets via TCP on port 23 from the host with DNS name A_SIGGEN1 as well as packets via TCP on port 23 on the host with DNS name A_SPECAN2. |
| “192.168.0.1” | Accepts trigger source packets from via TCP on the IANA registered port for LXI events (lxi-eventsvc) from the host at IP address 192.168.0.1. |
| “All:23,A_SPECAN2” | Accepts trigger source packets via UDP multicast on the IANA assigned multicast address to port 23 from any host as well as packets via TCP on the IANA registered port for LXI events (lxi-eventsvc) from the host with DNS name A_SPECAN2. |
| “” | Accepts trigger source packets via TCP or UDP multicast on the IANA assigned multicast address on the IANA registered port for LXI events (lxi-eventsvc) from any host. This is equivalent to <any> with the IANA registered port for LXI events (lxi-evntsvc). |
| “All” | Accepts trigger source packets via UDP multicast on the IANA assigned multicast address on the IANA registered port for LXI events (lxi-eventsvc) from any host. This is equivalent to “”. |
| “All:8543” | Accepts trigger source packets arriving via UDP multicast on the IANA UDP multicast address on port 8543 from any host. |
| “All, 192.168.1.1” | Accepts trigger source packets arriving via UDP multicast on the IANA assigned multicast address at the IANA registered port for LXI events (lxi-eventsvc), or TCP packets arriving on the IANA registered port for LXI events (lxi-eventsvc) from IP address 192.168.1.1 |
| “:23” | Accepts trigger source packets arriving via either UDP multicast or TCP on port 23 from any host. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

## IviLxiSyncTrigger Functions

The IviLxiSyncTrigger subsystem defines the following functions:

- Add Trigger Alarm

- Add Trigger Source

- Configure Trigger Alarm

- Configure Trigger Source

- Disable All Trigger Alarms

- Get Trigger Alarm Name (IVI-C Only)

- Get Trigger Source Name (IVI-C Only)

- Remove Trigger Alarm

- Remove Trigger Source

- Remove All Custom Trigger Sources

- Remove All Trigger Alarms

This section describes the behavior and requirements of each function.

    - Add Trigger Alarm

Description

This function creates a new trigger alarm.

When a new alarm is added, the default values for the associated attributes are as follows:

| Attribute | Default Value |
| --- | --- |
| Trigger Alarm Enabled | False |
| Trigger Alarm Period | 0 |
| Trigger Alarm Repeat Count | 1 |
| Trigger Alarm Time (.NET only) | 00:00:00 Jan 1, 1970 |
| Trigger Alarm Time Fraction (C/COM only) | 0 |
| Trigger Alarm Time Seconds (C/COM only) | 0 |

.NET Method Prototype

IIviLxiSyncTriggerAlarm Trigger.Alarms.Add(String alarmName);

COM Method Prototype

HRESULT Trigger.Alarms.Add([in] BSTR AlarmName);

C Prototype

ViStatus IviLxiSyncTriger_AddTriggerAlarm (ViSession Vi, 
                                           ViConstString AlarmName);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| AlarmName | Specifies the name of the trigger alarm to create. | ViConstString |
| Outputs | Description | Base Type |
| Return value (.NET) | A reference to the trigger alarm object that was added . | IIviLxiSyncTriggerAlarm |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Alarm Exists | Error: Alarm already exists. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

The table below specifies additional class-defined exceptions for this method.

| Exception Class | Description |
| --- | --- |
| AlarmExistsException | Alarm already exists. |

    - Add Trigger Source

Description

This function creates a new trigger source.

When a new alarm is added, the default values for the associated attributes are as follows:

| Attribute | Default Value |
| --- | --- |
| Trigger Source Delay | 0 |
| Trigger Source Detection | Positive |
| Trigger EventId | The repeated capability source name specified in this call |
| Trigger Source Filter | “” (Empty string) |

.NET Method Prototype

IIviLxiSyncTriggerSource Trigger.Sources.Add(String sourceName);

COM Method Prototype

HRESULT Trigger.Sources.Add([in] BSTR SourceName);

C Prototype

ViStatus IviLxiSyncTriger_AddTriggerSource (ViSession Vi, 
                                            ViConstString SourceName);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| SourceName | Specifies the name of the trigger source to create. | ViConstString |
| Outputs | Description | Base Type |
| Return value (.NET) | A reference to the trigger source object that was added . | IIviLxiSyncTriggerSource |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Event Source Exists | Error: Event source already exists. |
| Out of Event Resources | Error: Out of event resources. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

The table below specifies additional class-defined exceptions for this method.

| Exception Class | Description |
| --- | --- |
| EventSourceExistsException | Event source already exists. |
| OutOfEventResourcesException | Out of event resources. |

    - Configure Trigger Alarm

Description

This function configures the most commonly used attributes of the trigger alarm subsystem.

.NET Method Prototype

void Trigger.Alarms[].Configure(PrecisionDateTime time,
                                PrecisionTimeSpan period, 
                                Int32 repeatCount);

COM Method Prototype

HRESULT Trigger.Alarms.Item().Configure([in] DOUBLE TimeSeconds,
                                        [in] DOUBLE TimeFraction, 
                                        [in] DOUBLE Period, 
                                        [in] LONG RepeatCount);

C Prototype

ViStatus IviLxiSync_ConfigureTriggerAlarm (ViSession Vi, 
                                           ViConstString AlarmName, 
                                           ViReal64 TimeSeconds, 
                                           ViReal64 TimeFraction, 
                                           ViReal64 Period, 
                                           ViInt32 RepeatCount);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| AlarmName | The name of the alarm. | ViConstString |
| time (.NET) | Specifies the date and time of the trigger alarm. The driver uses this value to set the Trigger Alarm Time attribute. See the attribute description for more information. | Ivi.Driver. PrecisionDateTime |
| TimeSeconds (C/COM) | Specifies the seconds part of time. The driver uses this value to set the Trigger Alarm Time Seconds attribute. See the attribute description for more information. | ViReal64 |
| TimeFraction (C/COM) | Specifies the fractional part of time. The driver uses this value to set the Trigger Alarm Time Fractional attribute. See the attribute description for more information. | ViReal64 |
| Period | Specifies the period of the trigger alarm. The driver uses this value to set the Trigger Alarm Period attribute. See the attribute description for more information. | ViReal64 (C/COM) Ivi.Driver. PrecisionTimeSpan (.NET) |
| RepeatCount | Specifies the number of times to repeat the trigger at the period specified by the Trigger Alarm Repeat Period attribute. The driver uses this value to set the Trigger Alarm Repeat Count attribute. See the attribute description for more information. | ViInt32 |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Configure Trigger Source

Description

This function configures the most commonly used attributes of the trigger source subsystem.

.NET Method Prototype

void Trigger.Sources[].Configure(PrecisionTimeSpan delay,

Slope detection);

COM Method Prototype

HRESULT Trigger.Sources.Item().Configure([in] DOUBLE Delay,
                                         [in] IviLxiSyncSourceDetectionEnum Slope);

C Prototype

ViStatus IviLxiSync_ConfigureTriggerSource (ViSession Vi, 
                                            ViConstString SourceName, 
                                            ViReal64 Delay, 
                                            ViInt32 Detection);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| SourceName | Name of the trigger source to configure. | ViConstString |
| Delay | Specifies the trigger source delay. For C and COM, the units are seconds. For .NET, the units are implicit in the definition of PrecisionTimeSpan. A negative value implies pre-trigger acquisition. The driver uses this value to set the Trigger Source Delay attribute. See the attribute description for more information. | ViReal64 (C/COM) Ivi.Driver. PrecisionTimeSpan (.NET) |
| Detection | Specifies the slope of the trigger source. The driver uses this value to set the Trigger Source Detection attribute. See the attribute description for more information. | ViInt32 |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Disable All Trigger Alarms

Description

This function disables all trigger alarms.  The specific driver uses this function to set the Trigger Alarm Enabled property to False for all trigger alarms.

.NET Method Prototype

void Trigger.Alarms.DisableAll();

COM Method Prototype

HRESULT Trigger.Alarms.DisableAll();

C Prototype

ViStatus IviLxiSync_DisableAllTriggerAlarms (ViSession Vi);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Get Trigger Alarm Name (IVI-C Only)

Description

This function returns the physical repeated capability identifier that corresponds to the one-based index that the user specifies. If the value that the user passes for the AlarmIndex parameter is less than one or greater than the value of the Trigger Alarm Count attribute, the function returns an empty string in the AlarmName parameter and returns an error.  For custom trigger sources added vith the Add Arm Source function, this function returns the arm source name in the original casing used when Add Arm Source was called.

.NET Method Prototype

N/A

(use the Trigger.Alarms[].Name property)

COM Method Prototype

N/A

(Use the Trigger.Alarms.Item().Name property.)

C Prototype

ViStatus IviLxiSync_GetTriggerAlarmName (ViSession Vi, 
                                         ViInt32 AlarmIndex, 
                                         ViInt32 AlarmNameBufferSize, 
                                         ViChar AlarmName[]);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| AlarmIndex | A one-based index that defines which name to return. | ViInt32 |
| AlarmNameBufferSize | The number of bytes in the ViChar array that the user specifies for the AlarmName parameter. | ViInt32 |
| Outputs | Description | Data Type |
| AlarmName | The buffer into which the function returns the alarm name that corresponds to the index the user specifies. The caller may pass VI_NULL for this parameter if the AlarmNameBufferSize parameter is 0. | ViChar[] |

Return Values (C)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

    - Get Trigger Source Name (IVI-C Only)

Description

This function returns the physical repeated capability identifier that corresponds to the one-based index that the user specifies. If the value that the user passes for the SourceIndex parameter is less than one or greater than the value of the Trigger  Source Count attribute, the function returns an empty string in the SourceName parameter and returns an error.  For custom trigger sources added vith the Add Trigger Source function, this function returns the trigger source name in the original casing used when Add Trigger Source was called.

.NET Method Prototype

N/A

(use the Trigger.Sources[].Name property)

COM Method Prototype

N/A

(Use the Trigger.Sources.Item().Name property.)

C Prototype

ViStatus IviLxiSync_GetTriggerSourceName (ViSession Vi, 
                                          ViInt32 SourceIndex, 
                                          ViInt32 SourceNameBufferSize, 
                                          ViChar SourceName[]);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| SourceIndex | A one-based index that defines which name to return. | ViInt32 |
| SourceNameBufferSize | The number of bytes in the ViChar array that the user specifies for the SourceName parameter. | ViInt32 |
| Outputs | Description | Data Type |
| SourceName | The buffer into which the function returns the source name that corresponds to the index the user specifies. The caller may pass VI_NULL for this parameter if the SourceNameBufferSize parameter is 0. | ViChar[] |

Return Values (C)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

    - Remove Trigger Alarm

Description

This function removes a trigger alarm.

The AlarmName parameter is case-insensitive.

.NET Method Prototype

void Trigger.Alarms.Remove(String alarmName);

COM Method Prototype

HRESULT Trigger.Alarms.Remove([in] BSTR AlarmName);

C Prototype

ViStatus IviLxiSync_RemoveTriggerAlarm (ViSession Vi, 
                                        ViConstString AlarmName);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| AlarmName | Specifies the name of the trigger alarm to remove. | ViConstString |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Alarm Does Not Exist | Error: Alarm does not exist. |
| Attempt To Remove Reserved Repeated Capability | Error: The repeated capability name is reserved and cannot be removed. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

The table below specifies additional class-defined exceptions for this method.

| Exception Class | Description |
| --- | --- |
| AlarmDoesNotExistException | Alarm does not exist. |
| CannotRemoveReservedRepeatedCapabilityException | The repeated capability name is reserved and cannot be removed. |

    - Remove Trigger Source

Description

This function removes a trigger source.

The SourceName parameter is case-insensitive.

.NET Method Prototype

void Trigger.Sources.Remove(String sourceName);

COM Method Prototype

HRESULT Trigger.Sources.Remove([in] BSTR SourceName);

C Prototype

ViStatus IviLxiSync_RemoveTriggerSource (ViSession Vi, 
                                         ViConstString SourceName);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| SourceName | Specifies the name of the trigger source to remove. | ViConstString |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Event Source Does Not Exist | Error: Event source does not exist. |
| Attempt To Remove Reserved Repeated Capability | Error: The repeated capability name is reserved and cannot be removed. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

The table below specifies additional class-defined exceptions for this method.

| Exception Class | Description |
| --- | --- |
| EventSourceDoesNotExistException | Event source does not exist. |
| CannotRemoveReservedRepeatedCapabilityException | The repeated capability name is reserved and cannot be removed. |

    - Remove All Custom Trigger Sources

Description

This function removes all of the custom trigger sources that were added using the Add Trigger Source function.  The trigger sources associated with the reserved repeated capability identifiers, as defined in Section 2.1.6, Reserved Repeated Capability Identifiers, are not affected by this function.

.NET Method Prototype

void Trigger.Sources.RemoveAllCustomTriggerSources();

COM Method Prototype

HRESULT Trigger.Sources.RemoveAllCustomTriggerSources();

C Prototype

ViStatus IviLxiSync_RemoveAllCustomTriggerSources (ViSession Vi);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Remove All Trigger Alarms

Description

This function removes all of the trigger alarms that were added using the Add Trigger Alarm function.

.NET Method Prototype

void Trigger.Alarms.RemoveAllTriggerAlarms();

COM Method Prototype

HRESULT Trigger.Alarms.RemoveAllTriggerAlarms();

C Prototype

ViStatus IviLxiSync_RemoveAllTriggerAlarms (ViSession Vi);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

# IviLxiSyncEvent Subsystem

The IIviLxiSyncEvents interface provides a mechanism for configuration the generation of outbound events from an LXI device.  Events can be transmitted over the LXI trigger bus, as LXI-defined LAN events, or as custom user-defined events.  In addition, events can be transmitted to specific ports on specific hosts or to multiple hosts.

## Behavior Model

The diagram below shows the behavior model for the IviLxiSyncEvent subsystem.  This subsystem is responsible for routing signals to the appropriate event transmitter (either an LXI trigger bus line or a LAN event packet).  All signals (not just signals from the Arm-Trigger state machine) which are intended  to be utilized for sending events or routed to the LXI trigger bus, need to be connected to the input multiplexers in the event logic shown below.

Figure 5.  IviLxiSyncEvent Behavior Model

The LXI trigger bus may be used in either wired-OR or driven mode.  To use the bus in wired-OR mode, one LXI device must be configured to provide the wired-OR bias.  The device providing the bias can also either source events or respond to them on that trigger line.

For an LXI device to source events on the trigger bus, it must be configured to either drive the line in wired-OR or driven mode.  Therefore, when an event is configured to drive a trigger line, the mode of that line must be specified.

When an LXI device is configured to source events used the LXI LAN protocol, it must be configured to source LAN packets on the leading edge of the event, the falling edge of the event or both.  To accomplish this, if the event is configured to be wired-OR, then only logical true transitions are emitted.   If the event is not configured to be wired-OR, then both rising and falling transitions result in LAN packets being sourced.  Finally, if the system design requires only packets when the event becomes false, the device is configured in wired-OR mode, however the slope is set to false.

## IviLxiSyncEvent Attributes

The IviLxiSyncEvent subsystem defines the following attributes:

- Event Count

- Event Destination Path

- Event Drive Mode

- Event Item (IVI-COM and IVI.NET Only)

- Event Name (IVI-COM and IVI.NET Only)

- Event Slope

- Event Source

- Event Wired OR Bias Mode

This section describes the behavior and requirements of each attribute. The actual value for each attribute ID is defined in Section 8, Attribute ID Definitions.

    - Event Count

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | RO | IviLxiSyncEvent | None | N/A |

.NET Property Name

Events.Count

This property is inherited from IIviRepeatedCapabilityCollection.

COM Property Name

Events.Count

C Constant Name

IVILXISYNC_ATTR_EVENT_COUNT

Description

Returns the number of defined events.  The count returned includes any of the supported reserved repeated capability names defined in Section 2.1.6, Reserved Repeated Capability Identifiers as well as any custom repeated capability identifiers.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Event Destination Path

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | R/W | IviLxiSyncEvent | None | Configure Event |

.NET Property Name

Events[].DestinationPath

COM Property Name

Events.Item().DestinationPath

C Constant Name

IVILXISYNC_ATTR_EVENT_DESTINATION_PATH

Description

Specifies a list of places to send the event.

The default value for this attribute is the repeated capability name.

The grammar for the parameter is:

<DestinationPath>== [(<tcp>|<udp>|<TriggerBus>)[, <DestinationPath>]]

<tcp> == <host> [: <port>] [/<LANIdent>]

<udp> == [ALL] [: <port> ] [/<LANIdent>]

<TriggerBus> == LXI0|LXI1|LXI2|LXI3|LXI4|LXI5|LXI6|LXI7

host is either a hostname or host number, and port is a series of decimal digits indicating the port number.  Note that the hostname can not be “ALL” or one of the <TriggerBus> designations since that would indicate the <udp> or <TriggerBus> construct.

<LANIdent> is a string indicating the LAN identifier that will be sent in the LAN message.  The <LANIdent> is not case sensitive.  <LANIdent> is from one to 16 ASCII characters inclusive.  The characters may be numeric or underscore or hyphen or upper or lower-case alphabetic characters.

Defaults:

The default <DestinationPath> is the repeated capability name.  This may either be a <TriggerBus> identifier or a <LANIdent> identifier.

<LANIdent> defaults to the repeated capability name.

<port> defaults to the IANA registered port for LXI events (lxi-evntsvc).

If the repeated capability name is not a trigger bus specifier then the default <DestinationPath> is ‘ALL’ with the <LANIdent> as the repeated capability name.

If the repeated capability name corresponds to a <TriggerBus>, the default <DestinationPath> is the repeated capability name.

If multiple <DestinationPath>s are specified, the event is transmitted to each.

The <tcp> construct specifies that a TCP message will be sent to the destination when the bound event occurs.

The <udp> construct specifies that a UDP multicast message will be sent to the IANA registered multi-cast address (LXI-EVENT) on the designated port.  UDP unicast and UDP broadcasts are not supported by this syntax.

The <TriggerBus> construct specifies that a physical LXI wired trigger bus is used to signal the event.

Note that the LXI specification reserves event identifiers that begin with the characters “LXI” for LXI use.  The strings “LXI0”, “LXI1”, … ,”LXI7” refer to the 8 LXI wired trigger bus triggers.  See LXI 1.1 rule 6.4.5.

White space shall be ignored.  The <Destination> string is case insensitive.

Drivers may accept additional vendor-defined syntaxes

Drivers (and the corresponding instruments) that support this syntax are permitted to not support all possible destination syntaxes.

Conventional devices should consider restricting the port to only the IANA registered port for LXI events (lxi-evntsvc) and not accepting the generalized syntax.

| Example Destination Path | Description |
| --- | --- |
| “192.168.0.1:23/LAN2” | When the event designated by the event source property occurs, a TCP packet with event ID LAN2 is sent to 192.168.0.1, port 23. . |
| “LXI5” | When the event designated by the event source property occurs, the LXI5 wired-trigger-bus line generates a transition. |
| “ALL:23” | When the event designated by the event source property occurs, a UDP multicast packet is sent to the IANA assigned UDP multicast address port 23 with the event ID corresponding to the repeated capability name. |
| “A_SIGGEN1:23,A_SPECAN2:23” | When the event designated by the event source property occurs, TCP packets with the event ID corresponding to the repeated capability name are sent to hosts A_SIGGEN1 and A_SPECAN2 on port 23. |
| “192.168.0.1” | When the event designated by the event source property occurs, a TCP packet corresponding to the repeated capability name is sent to 192.168.0.1 on the IANA registered port for LXI events (lxi-eventsvc). |
| “:23” | When the event designated by the event source property occurs, a UDP multicast packet with the event id corresponding to the repeated capability name is sent to the IANA designated multicast address, port 23 |
| “:23,A_SPECAN2” | When the event designated by the event source property occurs, a UDP multicast packet is sent to the IANA designated multicast address, port 23, subsequently, a TCP packet is sent to host A_SPECAN2 on the IANA registered port for LXI events (lxi-eventsvc) |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Event Drive Mode

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | R/W | IviLxiSyncEvent | None | Configure Event |

.NET Property Name

Events[].DriveMode

COM Property Name

Events.Item().DriveMode

COM Enumeration Name

IviLxiSyncEventDriveModeEnum

C Constant Name

IVILXISYNC_ATTR_EVENT_DRIVE_MODE

Description

Specifies how this event is transmitted

It is an error to turn on the Wired OR Bias Mode for this device for a particular LXI trigger line and then set the Event Enabled attribute to On instead of Wired OR for an event whose destination path includes that LXI trigger line.

Defined Values

| Name | Description |  |  |
| --- | --- | --- | --- |
|  |  | Language | Identifier |
| On | Enables the event in driven mode. |  |  |
|  |  | .NET | EventDriveMode.On |
|  |  | C | IVILXISYNC_VAL_EVENT_ON |
|  |  | COM | IviLxiSyncEventDriveModeOn |
| Off | Disables the event. |  |  |
|  |  | .NET | EventDriveMode.Off |
|  |  | C | IVILXISYNC_VAL_EVENT_OFF |
|  |  | COM | IviLxiSyncEventDriveModeOff |
| Wired OR | Enables the event in wired-OR mode. |  |  |
|  |  | .NET | EventDriveMode.WiredOr |
|  |  | C | IVILXISYNC_VAL_EVENT_WIREDOR |
|  |  | COM | IviLxiSyncEventDriveModeWiredOr |

Return Values

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Event Source Not Set | Error: Event source not set. |
| Wired OR Mode Invalid | Error: Event source cannot operate in driven mode while serving as the wired-OR bias. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.  The table below specifies additional exceptionss for this property.

| Completion Codes | Description |
| --- | --- |
| Event Source Not Set Exception | Event source not set. |
| Wired OR Mode Invalid Exception | Event source cannot operate in driven mode while serving as the wired-OR bias. |

    - Event Item (IVI-COM and IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| IIviLxiSyncEvent* | RO | IviLxiSyncEvent | None | N/A |

.NET Property Name

Events[String name]

This indexer is inherited from the base interface IIviRepeatedCapabilityCollection.  The name parameter uniquely identifies a particular event in the events collection.

COM Property Name

Events.Item ([in] EventName)

C Constant Name

N/A

Description

Event Item uniquely identifies an event in the events collection. It returns an interface pointer which can be used to control the attributes and other functionality of that event.

The Item property may only take an event name.  If the user passes an invalid value for the event name parameter, the property returns an error.

Valid names include physical repeated capability identifiers and virtual repeated capability identifiers.

Return Values

If the IVI-COM driver cannot recognize the Name parameter, it returns an Unknown Name in Selector completion code as described in IVI-3.2: Inherent Capabilities Specification, Section 9.3.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Event Name (IVI-COM  and IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | RO | IviLxiSyncEvent | None | N/A |

.NET Property Name

Events[].Name

The .NET property is inherited from IIviRepeatedCapabilityIdentification.

COM Property Name

Events.Name ([in] LONG EventIndex)

C Constant Name

N/A

(Use the GetEventName function.)

Description

Returns the physical repeated capability identifier defined by the specific driver for the event that corresponds to the one-based index that the user specifies.  For custom event sources added vith the Add Event Source function, this function returns the event source name in the original casing used when Add Event Source was called.

For C and COM, valid values for the EventIndex parameter are between one and the value of the Event Count attribute. If the user passes an invalid value for the EventIndex parameter, the value of this attribute is an empty string.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Event Slope

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | R/W | IviLxiSyncEvent | None | Configure Event |

.NET Property Name

Events[].Slope

.NET Enumeration Name

Slope

COM Property Name

Events.Item().Slope

COM Enumeration Name

IviLxiSyncSourceSlopeEnum

C Constant Name

IVILXISYNC_ATTR_EVENT_SLOPE

Description

Specifies the slope of the event that is inbound to the event subsystem that will cause the generation of an outbound event.  The outbound event shall be transmitted with the same slope as the inbound event.

Defined Values

| .NET Exceptions Name | Description |  |  |
| --- | --- | --- | --- |
|  |  | Language | Identifier |
| Positive | The event will be transmitted with a rising edge. |  |  |
|  |  | .NET | Slope.Positive |
|  |  | C | IVILXISYNC_VAL_SLOPE_POSITIVE |
|  |  | COM | IviLxiSyncSourceSlopePositive |
| Negative | The event will be transmitted with a falling edge. |  |  |
|  |  | .NET | Slope.Negative |
|  |  | C | IVILXISYNC_VAL_SLOPE_NEGATIVE |
|  |  | COM | IviLxiSyncSourceSlopeNegative |

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Event Source

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | R/W | IviLxiSyncEvent | None | Configure Event |

.NET Property Name

Events[].Source

COM Property Name

Events.Item().Source

C Constant Name

IVILXISYNC_ATTR_EVENT_SOURCE

Description

Specifies the signal which causes an event to be transmitted.

In general, drivers will define events that are relevant for the instrument they support.

The following table lists reserved event source names that are defined by LXI and may be recognized by specific drivers according to their LXI Functional class.

| Table 5-3. Reserved Event Source Names |  |  |
| --- | --- | --- |
| Repeated Capability Identifier | Defined for LXI Functional Class |  |
| OperationComplete | Class A, Class B |  |
| Measuring | Class A, Class B |  |
| Settling | Class A, Class B |  |
| Sweeping | Class A, Class B |  |
| WaitingForArm | Class A, Class B |  |
| WaitingForTrigger | Class A, Class B |  |
| LAN0 | Class A, Class B |  |
| LAN1 | Class A, Class B |  |
| LAN2 | Class A, Class B |  |
| LAN3 | Class A, Class B |  |
| LAN4 | Class A, Class B |  |
| LAN5 | Class A, Class B |  |
| LAN6 | Class A, Class B |  |
| LAN7 | Class A, Class B |  |
| LXI0 | Class A |  |
| LXI1 | Class A |  |
| LXI2 | Class A |  |
| LXI3 | Class A |  |
| LXI3 | Class A |  |
| LXI5 | Class A |  |
| LXI6 | Class A |  |
| LXI7 | Class A |  |

Return Values

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Invalid Event Source | Error: Invalid source. |
| Out of Event Resources | Error: Out of event resources. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property. The table below specifies additional status codes for this property.

| Completion Codes | Description |
| --- | --- |
| Invalid Event Source Exception | Invalid source. |
| Out of Event Resources Exception | Out of event resources. |

    - Event Wired OR Bias Mode

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 (C/COM) | R/W | IviLxiSyncEvent | None | N/A |
| WiredOrBiasModeLines (.NET) | R/W | IviLxiSyncEvent | None | N/A |

.NET Property Name

Events.WiredOrBiasMode

.NET Enumeration Name

WiredOrBiasModeLines

COM Property Name

Events.WiredOrBiasMode

C Constant Name

IVILXISYNC_ATTR_EVENT_WIRED_OR_BIAS_MODE

Description

Specifies whether this LXI device will serve as the wired-OR bias for specific LXI trigger bus lines.

For C and COM, the allowed values for this attribute are 0 to 255.  This attribute is a bit field, where bit 0 represents LXI0, bit 1 represents LXI1, and so on.  A value of one in a particular bit indicates that the LXI device shall serve as the bias for the corresponding trigger bus line.  A value of zero in a particular bit disables the bias for the corresponding trigger bus line.  To use a trigger bus line in driven mode, the bias must be disabled.

For .NET, multiple lines are specified by bit-wise OR’ing enumeration values, since the enumeration is a flags enumeration.

Enabling wired-OR bias has no impact on the device’s ability to either respond to signals on trigger bus lines or to send events on trigger bus lines.

One and only one LXI device can serve as the wired-OR bias for a particular trigger bus line, although different devices can serve as the wired-OR bias for different trigger bus lines.

Defined Values

Note that the .NET enumeration is a flags enumeration.

| Return Values Name | Description |  |  |
| --- | --- | --- | --- |
|  |  | Language | Identifier |
| LXI0 | Wired Trigger Bus line LXI0. |  |  |
|  |  | .NET | WiredOrBiasModeLines.Lxi0 |
| LXI1 | Wired Trigger Bus line LXI1. |  |  |
|  |  | .NET | WiredOrBiasModeLines.Lxi1 |
| LXI2 | Wired Trigger Bus line LXI2. |  |  |
|  |  | .NET | WiredOrBiasModeLines.Lxi2 |
| LXI3 | Wired Trigger Bus line LXI3. |  |  |
|  |  | .NET | WiredOrBiasModeLines.Lxi3 |
| LXI4 | Wired Trigger Bus line LXI4. |  |  |
|  |  | .NET | WiredOrBiasModeLines.Lxi4 |
| LXI5 | Wired Trigger Bus line LXI5. |  |  |
|  |  | .NET | WiredOrBiasModeLines.Lxi5 |
| LXI6 | Wired Trigger Bus line LXI6. |  |  |
|  |  | .NET | WiredOrBiasModeLines.Lxi6 |
| LXI7 | Wired Trigger Bus line LXI7. |  |  |
|  |  | .NET | WiredOrBiasModeLines.Lxi7 |

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

## IviLxiSyncEvent Functions

The IviLxiSyncEvent subsystem defines the following functions:

- Add Event

- Configure Event

- Disable All Events

- GetEvent name (IVI-C Only)

- Remove Event

- Remove All Custom Events

This section describes the behavior and requirements of each function.

    - Add Event

Description

This function creates a new event.

When a new event is added, the default values for the associated attributes are as follows:

| Attribute | Default Value |
| --- | --- |
| Event Enabled | False |
| Event Slope | Positive |
| Event Source | “” (Empty string) |
| Event Destination Path | Current repeated capability instance name |

.NET Method Prototype

IIviLxiSyncEvent Events.Add(String eventName);

COM Method Prototype

HRESULT Events.Add([in] BSTR EventName);

C Prototype

ViStatus IviLxiSync_AddEvent (ViSession Vi, 
                              ViConstString EventName);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| EventName | Specifies the name of the event to create. | ViConstString |
| Outputs | Description | Base Type |
| Return value (.NET) | A reference to the event object that was added . | IIviLxiSyncEvent |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Event Source Exists | Error: Event source already exists. |
| Out of Event Resources | Error: Out of event resources. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

The table below specifies additional class-defined exceptions for this method.

| Exception Class | Description |
| --- | --- |
| EventSourceExistsException | Event source already exists. |
| OutOfEventResourcesException | Out of event resources. |

    - Configure Event

Description

This function configures the most commonly used attributes of the event subsystem.

.NET Method Prototype

void Events[].Configure(EventDriveMode driveMode,
                        String source,
                        String destinationPath, 
                        Slope slope);

COM Method Prototype

HRESULT Events.Item().Configure([in] IviLxiSyncEventDriveModeEnum DriveMode,
                                [in] BSTR Source,
                                [in] BSTR DestinationPath, 
                                [in] IviLxiSyncSourceSlopeEnum Slope);

C Prototype

ViStatus IviLxiSync_ConfigureEvent (ViSession Vi, 
                                    ViConstString repCapIdentifier, 
                                    ViInt32 DriveMode,
                                    ViConstString Source, 
                                    ViConstString DestinationPath, 
                                    ViInt32 Slope);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| repCapIdentifier | Specifies the name of the event to configure. | ViConstString |
| DriveMode | Specifies the mode of the event. The driver uses this value to set the Event Drive Mode attribute. See the attribute description for more information. | ViInt32 |
| Source | Specifies the signal which causes an event to be transmitted. The driver uses this value to set the Event Source attribute. See the attribute description for more information. | ViConstString |
| DestinationPath | Specifies a list of places to send the event. The driver uses this value to set the Event Destination Path attribute. See the attribute description for more information. | ViConstString |
| Slope | Specifies the slope of the event signal. The driver uses this value to set the Event Slope attribute. See the attribute description for more information. | ViInt32 |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Disable All Events

Description

This function disables all events.  The specific driver uses this function to set the Event Drive Mode property to Off  for all events.

.NET Method Prototype

void Events.DisableAll();

COM Method Prototype

HRESULT Events.DisableAll();

C Prototype

ViStatus IviLxiSync_DisableAllEvents (ViSession Vi);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Get Event Name (IVI-C Only)

Description

This function returns the physical repeated capability identifier that corresponds to the one-based index that the user specifies. If the value that the user passes for the EventIndex parameter is less than one or greater than the value of the Event Count attribute, the function returns an empty string in the EventName parameter and returns an error. For custom event sources added vith the Add Event Source function, this function returns the event source name in the original casing used when Add Event Source was called.

.NET Method Prototype

N/A

(Use the Events[].Name property)

COM Method Prototype

N/A

(Use the Events.Item().Name property)

C Prototype

ViStatus IviLxiSync_GetEventName (ViSession Vi, 
                                  ViInt32 EventIndex, 
                                  ViInt32 EventNameBufferSize, 
                                  ViChar EventName[]);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| EventIndex | A one-based index that defines which name to return. | ViInt32 |
| EventNameBufferSize | The number of bytes in the ViChar array that the user specifies for the EventName parameter. | ViInt32 |
| Outputs | Description | Data Type |
| EventName | The buffer into which the function returns the alarm name that corresponds to the index the user specifies. The caller may pass VI_NULL for this parameter if the EventNameBufferSize parameter is 0. | ViChar[] |

Return Values (C)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Remove Event

Description

This function removes an event.

The EventName parameter is case-insensitive.

.NET Method Prototype

void Events.Remove(String eventName);

COM Method Prototype

HRESULT Events.Remove([in] BSTR EventName);

C Prototype

ViStatus IviLxiSync_RemoveEvent (ViSession Vi, 
                                 ViConstString EventName);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| EventName | Specifies the name of the event to remove. | ViConstString |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.  The table below specifies additional status codes for this function.

| Completion Codes | Description |
| --- | --- |
| Event Source Does Not Exist | Error: Event source does not exist. |
| Attempt To Remove Reserved Repeated Capability | Error: The repeated capability name is reserved and cannot be removed. |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

The table below specifies additional class-defined exceptions for this method.

| Exception Class | Description |
| --- | --- |
| EventSourceDoesNotExistException | Event source does not exist. |
| CannotRemoveReservedRepeatedCapabilityException | The repeated capability name is reserved and cannot be removed. |

    - Remove All Custom Events

Description

This function removes all of the custom events that were added using the Add Event function.  The events associated with the reserved repeated capability identifiers, as defined in Section 2.1.6, Reserved Repeated Capability Identifiers, are not affected by this function.

.NET Method Prototype

void Events.RemoveAllCustomEvents();

COM Method Prototype

HRESULT Events.RemoveAllCustomEvents();

C Prototype

ViStatus IviLxiSync_RemoveAllCustomEvents (ViSession Vi);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

# IviLxiSyncEventLog Subsystem

## IviLxiSyncEventLog Attributes

The IviLxiSyncEventLog subsystem defines the following attributes:

- Event Log Enabled

- Event Log Entry Count

This section describes the behavior and requirements of each attribute. The actual value for each attribute ID is defined in Section 8, Attribute ID Definitions.

    - Event Log Entry Count

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | RO | N/A | None | N/A |

.NET Property Name

EventLog.EntryCount

COM Property Name

EventLog.EntryCount

C Constant Name

IVILXISYNC_ATTR_EVENT_LOG_ENTRY_COUNT

Description

Returns the number of event log entries available.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Event Log Enabled

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViBoolean | R/W | N/A | None | N/A |

.NET Property Name

EventLog.Enabled

COM Property Name

EventLog.Enabled

C Constant Name

IVILXISYNC_ATTR_EVENT_LOG_ENABLED

Description

If set to True, the LXI device enables the event logging feature.  If set to False, the LXI device disables the event logging feature.

Compliance Notes

The specific driver shall implement both the True and False values.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

## IviLxiSyncEventLog Functions

The IviLxiSyncEventLog subsystem defines the following functions:

- Clear Event Log

- Get Next Event Log Entry

This section describes the behavior and requirements of each function.

    - Clear Event Log Entries

Description

This function removes all existing entries from the event log.

.NET Method Prototype

void EventLog.ClearEntries();

COM Method Prototype

HRESULT EventLog.ClearEntries();

C Prototype

ViStatus IviLxiSync_ClearEventLogEntries (ViSession Vi);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

    - Get Next Event Log Entry

Description

This function retrieves and clears the oldest event log entry for the IVI session.  If there are no entries in the event log, then the function returns an empty string in the LogEntry parameter.

The following rules apply to the C interface of the Get Next Event Log Entry function:

- The function complies with the rules in IVI-3.2, Inherent Capabilities, Section 3.1.2.1, Additional Compliance Rules for C Functions with ViChar Array Output Parameters.

- If the user passes 0 for the LogEntryBufferSize parameter, the function does not clear the oldest event log entry from the list.

.NET Method Prototype

String EventLog.GetNextEntry();

COM Method Prototype

HRESULT EventLog.GetNextEntry([out, retval] BSTR* LogEntry);

C Prototype

ViStatus IviLxiSync_GetNextEventLogEntry (ViSession Vi, 
                                          ViInt32 LogEntryBufferSize,
                                          ViChar LogEntry[]);

Parameters

| Inputs | Description | Data Type |
| --- | --- | --- |
| Vi | Unique identifier for an IVI session. | ViSession |
| LogEntryBufferSize | The number of bytes in the ViChar array that the user specifies for the LogEntry parameter. | ViInt32 |
| Outputs | Description | Data Type |
| LogEntry (C/COM) | The buffer in which the function returns the oldest event log entry. Can be VI_NULL if LogEntryBufferSize is 0. | ViChar[] |
| Return Value (.NET) | The oldest event log entry. | String |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

# IviLxiSyncTime Subsystem

## IviLxiSyncTime Attributes

The IviLxiSyncTime subsystem defines the following attributes:

- Is Time Master

- Is Time Synchronized

- System Time (IVI.NET Only)

This section describes the behavior and requirements of each attribute. The actual value for each attribute ID is defined in Section 8, Attribute ID Definitions.

    - Is Time Master

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViBoolean | RO | N/A | None | N/A |

.NET Property Name

Time.IsMaster

COM Property Name

Time.IsMaster

C Constant Name

IVILXISYNC_ATTR_IS_TIME_MASTER

Description

If True, the device is the 1588 master.  If False, the device is not the 1588 master.

Compliance Notes

The specific driver shall implement both the True and False values.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - Is Time Synchronized

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViBoolean | RO | N/A | None | N/A |

.NET Property Name

Time.IsSynchronized

COM Property Name

Time.IsSynchronized

C Constant Name

IVILXISYNC_ATTR_IS_TIME_SYNCHRONIZED

Description

If True, the LXI device is synchronized with the 1588 master.  If False, the LXI device is not synchronized with the 1588 master.

Compliance Notes

The specific driver shall implement both the True and False values.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

    - System Time (IVI.NET Only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| Ivi.Driver.PrecisionDateTime | RO | N/A | None | N/A |

.NET Property Name

Time.SystemTime

COM Property Name

N/A

(Use the GetSystemTime method.)

C Constant Name

N/A

(Use the GetSystemTime function.)

Description

The system time, expressed as a PrecisionDateTime object.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

## IviLxiSyncTime Functions

The IviLxiSyncTime subsystem defines the following functions:

- Get System Time (IVI-C and IVI-COM Only)

This section describes the behavior and requirements of each function.

    - Get System Time (IVI-C and IVI-COM Only)

Description

This function retrieves the current 1588 time.

.NET Method Prototype

N/A

(Use the SystemTime property)

COM Method Prototype

HRESULT Time.GetSystemTime([in, out] double* TimeSeconds,
                           [in, out] double* TimeFractional);

C Prototype

ViStatus IviLxiSync_GetSystemTime (ViSession Vi, 
                                   ViReal64* TimeSeconds,
                                   ViReal64* TimeFractional);

Parameters

| Inputs | Description | Data Type |
| --- | --- | --- |
| Vi | Unique identifier for an IVI session. | ViSession |
| Outputs | Description | Data Type |
| TimeSeconds | Indicates the seconds portion of the current 1588 time. | ViReal64 |
| TimeFractional | Indicates the fractional portion of the current 1588 time. | ViReal64 |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

# Attribute ID Definitions

The following table defines the ID value for all IviLxiSync attributes.  Each value is defined in terms of the IVILXISYNC_ATTR_BASE base ID, which is itself defined in terms of IVI_ATTR_BASE as follows:

IVILXISYNC_ATTR_BASE = IVI_ATTR_BASE + 950000

| Table 8-4. IviLxiSync Attribute ID Values |  |
| --- | --- |
| Attribute Name | ID Value |
| IVILXISYNC_ATTR_ARM_ALARM_COUNT | IVILXISYNC_ATTR_BASE + 1 |
| IVILXISYNC_ATTR_ARM_ALARM_ENABLED | IVILXISYNC_ATTR_BASE + 2 |
| IVILXISYNC_ATTR_ARM_ALARM_PERIOD | IVILXISYNC_ATTR_BASE + 3 |
| IVILXISYNC_ATTR_ARM_ALARM_REPEAT_COUNT | IVILXISYNC_ATTR_BASE + 4 |
| IVILXISYNC_ATTR_ARM_ALARM_TIME_FRACTION | IVILXISYNC_ATTR_BASE + 5 |
| IVILXISYNC_ATTR_ARM_ALARM_TIME_SECONDS | IVILXISYNC_ATTR_BASE + 6 |
| IVILXISYNC_ATTR_ARM_COUNT | IVILXISYNC_ATTR_BASE + 7 |
| IVILXISYNC_ATTR_ARM_DELAY | IVILXISYNC_ATTR_BASE + 8 |
| IVILXISYNC_ATTR_ARM_SOURCE_COUNT | IVILXISYNC_ATTR_BASE + 9 |
| IVILXISYNC_ATTR_ARM_SOURCE_DETECTION | IVILXISYNC_ATTR_BASE + 10 |
| IVILXISYNC_ATTR_ARM_SOURCE_ENABLED | IVILXISYNC_ATTR_BASE + 11 |
| IVILXISYNC_ATTR_ARM_SOURCE_EVENTID | IVILXISYNC_ATTR_BASE + 12 |
| IVILXISYNC_ATTR_ARM_SOURCE_FILTER | IVILXISYNC_ATTR_BASE + 13 |
| IVILXISYNC_ATTR_ARM_SOURCE_OR_ENABLED | IVILXISYNC_ATTR_BASE + 14 |
| IVILXISYNC_ATTR_TRIGGER_ALARM_COUNT | IVILXISYNC_ATTR_BASE + 15 |
| IVILXISYNC_ATTR_TRIGGER_ALARM_ENABLED | IVILXISYNC_ATTR_BASE + 16 |
| IVILXISYNC_ATTR_TRIGGER_ALARM_PERIOD | IVILXISYNC_ATTR_BASE + 17 |
| IVILXISYNC_ATTR_TRIGGER_ALARM_REPEAT_COUNT | IVILXISYNC_ATTR_BASE + 18 |
| IVILXISYNC_ATTR_TRIGGER_ALARM_TIME_FRACTION | IVILXISYNC_ATTR_BASE + 19 |
| IVILXISYNC_ATTR_TRIGGER_ALARM_TIME_SECONDS | IVILXISYNC_ATTR_BASE + 20 |
| IVILXISYNC_ATTR_TRIGGER_COUNT | IVILXISYNC_ATTR_BASE + 21 |
| IVILXISYNC_ATTR_TRIGGER_SOURCE | IVILXISYNC_ATTR_BASE + 22 |
| IVILXISYNC_ATTR_TRIGGER_SOURCE_COUNT | IVILXISYNC_ATTR_BASE + 23 |
| IVILXISYNC_ATTR_TRIGGER_SOURCE_DELAY | IVILXISYNC_ATTR_BASE + 24 |
| IVILXISYNC_ATTR_TRIGGER_SOURCE_DETECTION | IVILXISYNC_ATTR_BASE + 25 |
| IVILXISYNC_ATTR_TRIGGER_SOURCE_EVENTID | IVILXISYNC_ATTR_BASE + 26 |
| IVILXISYNC_ATTR_TRIGGER_SOURCE_FILTER | IVILXISYNC_ATTR_BASE + 27 |
| IVILXISYNC_ATTR_EVENT_COUNT | IVILXISYNC_ATTR_BASE + 28 |
| IVILXISYNC_ATTR_EVENT_DESTINATION_PATH | IVILXISYNC_ATTR_BASE + 29 |
| IVILXISYNC_ATTR_EVENT_DRIVE_MODE | IVILXISYNC_ATTR_BASE + 30 |
| IVILXISYNC_ATTR_EVENT_SLOPE | IVILXISYNC_ATTR_BASE + 31 |
| IVILXISYNC_ATTR_EVENT_SOURCE | IVILXISYNC_ATTR_BASE + 32 |
| IVILXISYNC_ATTR_EVENT_WIRED_OR_BIAS_MODE | IVILXISYNC_ATTR_BASE + 33 |
| IVILXISYNC_ATTR_EVENT_LOG_ENABLED | IVILXISYNC_ATTR_BASE + 34 |
| IVILXISYNC_ATTR_EVENT_LOG_ENTRY_COUNT | IVILXISYNC_ATTR_BASE + 35 |
| IVILXISYNC_ATTR_IS_TIME_MASTER | IVILXISYNC_ATTR_BASE + 36 |
| IVILXISYNC_ATTR_IS_TIME_SYNCHRONIZED | IVILXISYNC_ATTR_BASE + 37 |

# Attribute Value Definitions

This section specifies the actual value for each defined attribute value.

Arm and Trigger Alarm Repeat Count

| Value Name | Language | Identifier | Actual Value |
| --- | --- | --- | --- |
| Continuous | C | IVILXISYNC_VAL_REPEAT_CONTINUOUS | 0 |
|  | COM | N/A | N/A |

Arm Source Detection

| Value Name | Language | Identifier | Actual Value |
| --- | --- | --- | --- |
| Rise | .NET | ArmSourceDetection.Rise | 0 |
|  | C | IVILXISYNC_VAL_DETECTION_RISE | 0 |
|  | COM | IviLxiSyncArmSourceDetectionRise | 0 |
| Fall | .NET | ArmSourceDetection.Fall | 1 |
|  | C | IVILXISYNC_VAL_DETECTION_FALL | 1 |
|  | COM | IviLxiSyncArmSourceDetectionFall | 1 |
| High | .NET | ArmSourceDetection.High | 2 |
|  | C | IVILXISYNC_VAL_DETECTION_HIGH | 2 |
|  | COM | IviLxiSyncArmSourceDetectionHigh | 2 |
| Low | .NET | ArmSourceDetection.Low | 3 |
|  | C | IVILXISYNC_VAL_DETECTION_LOW | 3 |
|  | COM | IviLxiSyncArmSourceDetectionLow | 3 |

Event Drive Mode

| Value Name | Language | Identifier | Actual Value |
| --- | --- | --- | --- |
| On | .NET | EventDriveMode.Driven | 0 |
|  | C | IVILXISYNC_VAL_EVENT_DRIVEN | 0 |
|  | COM | IviLxiSyncEventDriveModeDriven | 0 |
| Off | .NET | EventDriveMode.Off | 1 |
|  | C | IVILXISYNC_VAL_EVENT_OFF | 1 |
|  | COM | IviLxiSyncEventDriveModeOn | 1 |
| Wired OR | .NET | EventDriveMode.WiredOr | 2 |
|  | C | IVILXISYNC_VAL_EVENT_WIREDOR | 2 |
|  | COM | IviLxiSyncEventDriveModeWiredOr | 2 |

Source Slope

| Value Name | Language | Identifier | Actual Value |
| --- | --- | --- | --- |
| Rise | .NET | Slope.Positive | 0 |
|  | C | IVILXISYNC_VAL_SLOPE_RISE | 0 |
|  | COM | IviLxiSyncSourceSlopeRise | 0 |
| Fall | .NET | Slope.Negative | 1 |
|  | C | IVILXISYNC_VAL_SLOPE_FALL | 1 |
|  | COM | IviLxiSyncSourceSlopeFall | 1 |

Trigger Source Detection

| Value Name | Language | Identifier | Actual Value |
| --- | --- | --- | --- |
| Rise | .NET | Slope.Positive | 0 |
|  | C | IVILXISYNC_VAL_DETECTION_RISE | 0 |
|  | COM | IviLxiSyncTriggerSourceDetectionRise | 0 |
| Fall | .NET | Slope.Negative | 1 |
|  | C | IVILXISYNC_VAL_DETECTION_FALL | 1 |
|  | COM | IviLxiSyncTriggerSourceDetectionFall | 1 |

# Function Parameter Value Definitions

This section specifies the actual values for each function parameter that defines values.

Configure Arm Source

Legal values for the Detection parameter are the same as those for Arm Source Detection in section 9, Attribute Value Definitions.

Configure Event

Legal values for the DriveMode parameter are the same as those for Event Drive Mode in section 9, Attribute Value Definitions.

Legal values for the Slope parameter are the same as those for Source Slope in section 9, Attribute Value Definitions.

Configure Trigger Source

Legal values for the Detection parameter are the same as those for Trigger Source Detection in section 9, Attribute Value Definitions.

# Error and Completion Code Value Definitions

The table below specifies the actual value for each status code that the IviLxiSync specification defines.

| Table 11-1. IviLxiSync Error and Completion Codes |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Error Name | Description |  |  |  |  |
|  |  | Language |  | Identifier | Value(hex) |
| Alarm Time Invalid | The alarm time is not valid. For instance, the time may have already passed. |  |  |  |  |
|  |  | .NET | AlarmTimeInvalidException |  | N/A |
|  |  | C | IVILXISYNC_ALARM_TIME_INVALID |  | 0xBFFA3001 |
|  |  | COM | E_IVILXISYNC_ALARM_TIME_INVALID |  | 0x80043001 |
| Event Source Exists | The event source is already a member of the collection. |  |  |  |  |
|  |  | .NET | EventSourceExistsException |  | N/A |
|  |  | C | IVILXISYNC_EVENT_SOURCE_EXISTS |  | 0xBFFA3002 |
|  |  | COM | E_IVILXISYNC_EVENT_SOURCE_EXISTS |  | 0x80043002 |
| Out of Event Resources | The device has no more event resources to allocate. |  |  |  |  |
|  |  | .NET | OutOfEventResourcesException |  | N/A |
|  |  | C | IVILXISYNC_OUT_OF_EVENT_RESOURCES |  | 0xBFFA3003 |
|  |  | COM | E_IVILXISYNC_OUT_OF_EVENT_RESOURCES |  | 0x80043003 |
| Event Source Does Not Exist | The specified event source has not been defined. |  |  |  |  |
|  |  | .NET | EventSourceDoesNotExistException |  | N/A |
|  |  | C | IVILXISYNC_EVENT_SOURCE_DOES_NOT_EXISTS |  | 0xBFFA3004 |
|  |  | COM | E_IVILXISYNC_EVENT_SORUCE_DOES_NOT_EXIST |  | 0x80043004 |
| Event Source Not Set | The event cannot be enabled while the event source has not been set. |  |  |  |  |
|  |  | .NET | EventSourceNotSetException |  | N/A |
|  |  | C | IVILXISYNC_EVENT_SOURCE_NOT_SET |  | 0xBFFA3005 |
|  |  | COM | E_IVILXISYNC_EVENT_SOURCE_NOT_SET |  | 0x80043005 |
| Invalid Event Source | The specified name is not a valid event source. |  |  |  |  |
|  |  | .NET | InvalidEventSourceException |  | N/A |
|  |  | C | IVILXISYNC_INVALID_EVENT_SOURCE |  | 0xBFFA3006 |
|  |  | COM | E_IVILXISYNC_INVALID_EVENT_SOURCE |  | 0x80043006 |
| Alarm Exists | The alarm is already a member of the collection. |  |  |  |  |
|  |  | .NET | AlarmExistsException |  | N/A |
|  |  | C | IVILXISYNC_ALARM_EXISTS |  | 0xBFFA3007 |
|  |  | COM | E_IVILXISYNC_ALARM_EXISTS |  | 0x80043007 |
| Alarm Does Not Exist | The specified alarm has not been defined. |  |  |  |  |
|  |  | .NET | AlarmDoesNotExistException |  | N/A |
|  |  | C | IVILXISYNC_ALARM_DOES_NOT_EXIST |  | 0xBFFA3008 |
|  |  | COM | E_IVILXISYNC_ALARM_DOES_NOT_EXIST |  | 0x80043008 |
| Wired OR Mode Invalid | The event source cannot operate in driven mode while serving as the wired-OR bias. |  |  |  |  |
|  |  | .NET | WiredOrModeInvalidException |  | N/A |
|  |  | C | IVILXISYNC_WIRED_OR_MODE_INVALID |  | 0xBFFA3009 |
|  |  | COM | E_IVILXISYNC_WIRED_OR_MODE_INVALID |  | 0x80043009 |
| Attempt To Remove Reserved Repeated Capability | The repeated capability name is reserved and cannot be removed. |  |  |  |  |
|  |  | .NET | CannotRemoveReservedRepeatedCapabilityException |  | N/A |
|  |  | C | IVILXISYNC_CANT_REMOVE_RESERVED_REPEATED_CAPABILITY |  | 0xBFFA3010 |
|  |  | COM | E_IVILXISYNC_CANT_REMOVE_RESERVED_REPEATED_CAPABILITY |  | 0x80043010 |

Table 10-2 defines the recommended format of the message string associated with the errors. In C, these strings are returned by the Get Error function. In COM, these strings are the description contained in the ErrorInfo object.

Note: In the description string table entries listed below, %s is always used to represent the component name.

| Table 11-2. IviLxiSync Error Message Strings |  |
| --- | --- |
| Name | Message String |
| Alarm Time Invalid | “%s: The alarm time is invalid” |
| Event Source Exists | “%s: The event source already exists” |
| Out of Event Resources | “%s: Out of event resources” |
| Event Source Does Not Exist | “%s: The specified event source does not exist” |
| Event Source Not Set | “%s: The event source has not been specified” |
| Invalid Event Source | “%s: The specified event source is not valid” |
| Alarm Exists | “%s: The alarms already exists” |
| Alarm Does Not Exist | “%s: The specified alarm has not been defined” |
| Wired OR Mode Invalid | “%s: The event source cannot operate in driven mode while serving as the wired-OR bias” |
| Attempt to Remove Reserved Repeated Capability | “%s: The reserved repeated capability cannot be removed from the collection” |

## IVI.NET IviLxiSync Exceptions and Warnings

This section defines the list of IVI.NET exceptions and warnings that are specific to the IviLxiSync class. For general information on IVI.NET exceptions and warnings, refer to IVI-3.1: Driver Architecture Specification and section 12, Common IVI.NET Exceptions and Warnings, of IVI-3.2: Inherent Capabilities Specification.

The IVI.NET exceptions defined in this specification are declared in the Ivi.LxiSync namespace.

- AlarmDoesNotExistException

- AlarmExistsException

- AlarmTimeInvalidException

- CannotResolveReservedRepeatedCapabilityException

- EventSourceDoesNotExistException

- EventSourceExistsException

- EventSourceNotSetException

- InvalidEventSourceException

- OutOfEventResourcesException

- WiredOrModeInvalidException

    - AlarmDoesNotExistException

Description

This exception is used when the driver finds that a specified alarm has not been defined.

Constructors

Ivi.LxiSync.AlarmDoesNotExistException(String message

String alarmName);

Ivi.LxiSync.AlarmDoesNotExistException();

Ivi.LxiSync.AlarmDoesNotExistException(String message);

Ivi.LxiSync.AlarmDoesNotExistException(String message,

System.Exception innerException);

Message String

The specified alarm has not been defined.

Alarm name: <alarmName>

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| alarmName | The name of the alarm that is undefined. | String |

Usage

If driver developers use constructors that take a message string, they are responsible for message string localization.

    - AlarmExistsException

Description

This exception is used when the driver finds that a specified alarm already exists.

Constructors

Ivi.LxiSync.AlarmExistsException(String message

String alarmName);

Ivi.LxiSync.AlarmExistsException();

Ivi.LxiSync.AlarmExistsException(String message);

Ivi.LxiSync.AlarmExistsException(String message,

System.Exception innerException);

Message String

The specified alarm already exists.

Alarm name: <alarmName>

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| alarmName | The name of the alarm. | String |

Usage

If driver developers use constructors that take a message string, they are responsible for message string localization.

    - AlarmTimeInvalidException

Description

This exception is thrown when an alarm time is invalid.

Constructors

Ivi.LxiSync.AlarmTimeInvalidException(String alarmName,

String alarmTime);

Ivi.LxiSync.AlarmTimeInvalidException();

Ivi.LxiSync.AlarmTimeInvalidException(String message);

Ivi.LxiSync.AlarmTimeInvalidException(String message,

System.Exception innerException);

Message String

The alarm time is invalid.

Alarm name: <alarmName>.

Alarm time: <alarmTime>.

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| alarmName | The name of the alarm. | String |
| alarmTime | The invalid alarm time | String |

Usage

If driver developers use constructors that take a message string, they are responsible for message string localization.

    - CannotResolveReservedRepeatedCapabilityException

Description

This exception is thrown when a reserved repeated capability cannot be removed from one of the IviLxiSync collections.

Constructors

Ivi.LxiSync.CannotResolveReservedRepeatedCapabilityException(

String repeatedCapability,

String repeatedCapabilityInstance);

Ivi.LxiSync.CannotResolveReservedRepeatedCapabilityException();

Ivi.LxiSync.CannotResolveReservedRepeatedCapabilityException(

String message);

Ivi.LxiSync.CannotResolveReservedRepeatedCapabilityException(

String message,

System.Exception innerException);

Message String

The reserved repeated capability cannot be removed from the collection.

Repeated capability name: <repeatedCapability>.

Repeated capability instance: <repeatedCapabilityInstance>.

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| repeatedCapability | The name of the repeated capability. | String |
| repeatedCapabilityInstance | The name of the repeated capability instance that cannot be removed from its collection. | String |

Usage

If driver developers use constructors that take a message string, they are responsible for message string localization.

    - EventSourceDoesNotExistException

Description

This exception is used when the driver finds that a specified event source is not defined.

Constructors

Ivi.LxiSync.EventSourceDoesNotExistException(String message,

String eventSourceName);

Ivi.LxiSync.EventSourceDoesNotExistException();

Ivi.LxiSync.EventSourceDoesNotExistException(String message);

Ivi.LxiSync.EventSourceDoesNotExistException(String message,

System.Exception innerException);

Message String

The specified event source has not been defined.

Event source name: <eventSourceName>

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| eventSourceName | The name of the event source that is undefined. | String |

Usage

If driver developers use constructors that take a message string, they are responsible for message string localization.

    - EventSourceExistsException

Description

This exception is used when the driver finds that a specified event source already exists.

Constructors

Ivi.LxiSync.EventSourceExistsException(String message,

String eventSourceName);

Ivi.LxiSync.EventSourceExistsException();

Ivi.LxiSync.EventSourceExistsException(String message);

Ivi.LxiSync.EventSourceExistsException(String message,

System.Exception innerException);

Message String

The specified event source already exists.

Event source name: <eventSourceName>

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| eventSourceName | The name of the event source that already exists. | String |

Usage

If driver developers use constructors that take a message string, they are responsible for message string localization.

    - EventSourceNotSetException

Description

This exception is used when the driver finds that the event source has not been specified.

Constructors

Ivi.LxiSync.EventSourceNotSetException();

Ivi.LxiSync.EventSourceNotSetException(String message);

Ivi.LxiSync.EventSourceNotSetException(String message,

System.Exception innerException);

Message String

The event source has not been specified.

Usage

If driver developers use constructors that take a message string, they are responsible for message string localization.

    - InvalidEventSourceException

Description

This exception is used when the driver finds that a specified event source is not valid.

Constructors

Ivi.LxiSync.InvalidEventSourceException(String message,

String eventSourceName);

Ivi.LxiSync.InvalidEventSourceException();

Ivi.LxiSync.InvalidEventSourceException(String eventSourceName);

Ivi.LxiSync.InvalidEventSourceException(String eventSourceName,

System.Exception innerException);

Message String

The specified event source is not valid.

Event source name: <eventSourceName>

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| eventSourceName | The name of the invalid event source. | String |

Usage

If driver developers use constructors that take a message string, they are responsible for message string localization.

    - OutOfEventResourcesException

Description

This exception is used when the driver is out of event resources.

Constructors

Ivi.LxiSync.OutOfEventResourcesException();

Ivi.LxiSync.OutOfEventResourcesException(String message);

Ivi.LxiSync.OutOfEventResourcesException(String message,

System.Exception innerException);

Message String

Out of event resources.

Usage

If driver developers use constructors that take a message string, they are responsible for message string localization.

    - WiredOrModeInvalidException

Description

This exception is used when the event source cannot operate in driven mode while serving as the wired-OR bias.

Constructors

Ivi.LxiSync.WiredOrModeInvalidException();

Ivi.LxiSync.WiredOrModeInvalidException(String message);

Ivi.LxiSync.WiredOrModeInvalidException(String message,

System.Exception innerException);

Message String

The event source cannot operate in driven mode while serving as the wired-OR bias.

Usage

If driver developers use constructors that take a message string, they are responsible for message string localization.

# Hierarchies

## .NET Hierarchy

| Table 12-5. .NET Hierarchy |  |  |
| --- | --- | --- |
| .NET Interface Hierarchy | Generic Name | Type |
| Arm |  |  |
| ArmCount | Arm Count | P |
| Delay | Arm Delay | P |
| Sources |  |  |
| Add | Add Arm Source | M |
| DisableAll | Disable All Arm Sources | M |
| OrEnabled | Arm Source Or Enabled | P |
| Remove | Remove Arm Source | M |
| RemoveAllCustomArmSources | Remove All Custom Arm Sources | M |
| Count | Arm Source Count | P |
| [] | Item |  |
| Configure | Configure Arm Source | M |
| Detection | Arm Source Detection | P |
| Enabled | Arm Source Enabled | P |
| EventID | Arm Source Event ID | P |
| Filter | Arm Source Filter | P |
| Name | Arm Source Name | P |
| Alarms |  |  |
| Add | Add Arm Alarm | M |
| DisableAll | Disable All Arm Alarms | M |
| Remove | Remove Arm Alarm | M |
| RemoveAllArmAlarms | Remove All Arm Alarms | M |
| Count | Arm Alarm Count | P |
| [] | Item |  |
| Configure | Configure Arm Alarm | M |
| Enabled | Arm Alarm Enabled | P |
| Name | Arm Alarm Name | P |
| Period | Arm Alarm Period | P |
| RepeatCount | Arm Alarm Repeat Count | P |
| Time | Arm Alarm Time | P |
| Trigger |  |  |
| TriggerCount | Trigger Count | P |
| TriggerSource | Trigger Source | P |
| Alarms |  |  |
| Add | Add Trigger Alarm | M |
| DisableAll | Disable All Trigger Alarms | M |
| Remove | Remove Trigger Alarm | M |
| RemoveAllTriggerAlarms | Remove All Trigger Alarms | M |
| Count | Trigger Alarm Count | P |
| [] | Item |  |
| Configure | Configure Trigger Alarm | M |
| Enabled | Trigger Alarm Enabled | P |
| Name | Trigger Alarm Name | P |
| Period | Trigger Alarm Period | P |
| RepeatCount | Trigger Alarm Repeat Count | P |
| Time | Trigger Alarm Time | P |
| Sources |  |  |
| Add | Add Trigger Source | M |
| Remove | Remove Trigger Source | M |
| RemoveAllCustomTriggerSources | Remove All Custom Trigger Sources | M |
| Count | Trigger Source Count | P |
| [] | Item |  |
| Configure | Configure Trigger Source | M |
| Delay | Trigger Source Delay | P |
| Detection | Trigger Source Detection | P |
| EventID | Trigger Source Event ID | P |
| Filter | Trigger Source Filter | P |
| Name | Trigger Source Name | P |
| Events |  |  |
| Add | Add Event | M |
| DisableAll | Disable All Events | M |
| Remove | Remove Event | M |
| RemoveAllCustomEvents | Remove All Custom Events | M |
| Count | Event Count | P |
| WiredOrBiasMode | Wired OR Bias Mode | P |
| [] | Item |  |
| Configure | Configure Event | M |
| DestinationPath | Event Destination Path | P |
| DriveMode | Event Drive Mode | P |
| Name | Event Name | P |
| Source | Event Source | P |
| Slope | Event Slope | P |
| EventLog |  |  |
| Clear | Clear Event Log | M |
| Enabled | Event Log Enabled | P |
| EntryCount | Event Log Entry Count | P |
| GetNextEntry | Get Next Event Log Entry | M |
| Time |  |  |
| SystemTime | System Time | M |
| IsMaster | Is Master | P |
| IsSynchronized | Is Synchronized | P |

### IviLxiSync .NET Interfaces

IviLxiSync-interfaces contain interface reference properties for accessing the following IviLxiSync interfaces:

- IIviLixSyncArm

- IIviLixSyncEventLog

- IIviLixSyncEvents

- IIviLixSyncTime

- IIviLixSyncTrigger

The IIviLxiSyncArm interface contains interface reference properties for accessing additional the following IviLxiSync interfaces:

- IIviLxiSyncArmAlarms

- IIviLxiSyncArmSources

The IIviLxiSyncArmAlarms interface contains methods and properties for accessing a collection of objects that implement the IIviLxiSyncArmAlarm interface.

The IIviLxiSyncArmSources interface contains methods and properties for accessing a collection of objects that implement the IIviLxiSyncArmSource interface.

The IIviLxiSyncTrigger interface contains interface reference properties for accessing additional the following IviLxiSync interfaces:

- IIviLxiSyncTriggerAlarms

- IIviLxiSyncTriggerSources

The IIviLxiSyncTriggerAlarms interface contains methods and properties for accessing a collection of objects that implement the IIviLxiSyncTriggerAlarm interface.

The IIviLxiSyncTriggerSources interface contains methods and properties for accessing a collection of objects that implement the IIviLxiSyncTriggerSource interface.

The IIviLxiSyncEvents interface contains methods and properties for accessing a collection of objects that implement the IIviLxiSyncEvent interface.

### Interface Reference Properties

Interface reference properties are used to navigate the IviLxiSync .NET hierarchy. This section describes the interface reference properties that the IIviLxiSync, IIviLxiSyncArm, IIviLxiSyncArmAlarms, IIviLxiSyncArmSources, IIviLxiSyncTriggerAlarms, IIviLxiSyncTriggerSources, and IIviLxiSyncEvents  interfaces define.  All interface reference properties are read-only.

| Table 12-6. Interface Referece Properties |  |
| --- | --- |
| Data Type | Access |
| IIviLxiSyncArm | Arm |
| IIviLxiSyncArmAlarm | Alarms[] |
| IIviLxiSyncArmAlarms | Alarms |
| IIviLxiSyncArmSource | Sources[] |
| IIviLxiSyncArmSources | Sources |
| IIviLxiSyncEvent | Events[] |
| IIviLxiSyncEventLog | EventLog |
| IIviLxiSyncEvents | Events |
| IIviLxiSyncTime | Time |
| IIviLxiSyncTrigger | Trigger |
| IIviLxiSyncTriggerAlarm | Alarms[] |
| IIviLxiSyncTriggerAlarms | Alarms |
| IIviLxiSyncTriggerSource | Sources[] |
| IIviLxiSyncTriggerSources | Sources |

## COM Hierarchy

| Table 12-3. COM Hierarchy |  |  |
| --- | --- | --- |
| COM Interface Hierarchy | Generic Name | Type |
| Arm |  |  |
| ArmCount | Arm Count | P |
| Delay | Arm Delay | P |
| Sources |  |  |
| Add | Add Arm Source | M |
| DisableAll | Disable All Arm Sources | M |
| OrEnabled | Arm Source Or Enabled | P |
| Remove | Remove Arm Source | M |
| RemoveAllCustomArmSources | Remove All Custom Arm Sources | M |
| Count | Arm Source Count | P |
| Name | Arm Source Name | P |
| Item |  |  |
| Configure | Configure Arm Source | M |
| Detection | Arm Source Detection | P |
| Enabled | Arm Source Enabled | P |
| Event ID | Arm Source Event ID | P |
| Filter | Arm Source Filter | P |
| Alarms |  |  |
| Add | Add Arm Alarm | M |
| DisableAll | Disable All Arm Alarms | M |
| Remove | Remove Arm Alarm | M |
| RemoveAllArmAlarms | Remove All Arm Alarms | M |
| Count | Arm Alarm Count | P |
| Name | Arm Alarm Name | P |
| Item |  |  |
| Configure | Configure Arm Alarm | M |
| Enabled | Arm Alarm Enabled | P |
| Period | Arm Alarm Period | P |
| RepeatCount | Arm Alarm Repeat Count | P |
| TimeSeconds | Arm Alarm Time Seconds | P |
| TimeFraction | Arm Alarm Time Fraction | P |
| Trigger |  |  |
| TriggerCount | Trigger Count | P |
| TriggerSource | Trigger Source | P |
| Alarms |  |  |
| Add | Add Trigger Alarm | M |
| DisableAll | Disable All Trigger Alarms | M |
| Remove | Remove Trigger Alarm | M |
| RemoveAllTriggerAlarms | Remove All Trigger Alarms | M |
| Count | Trigger Alarm Count | P |
| Name | Trigger Alarm Name | P |
| Item |  |  |
| Configure | Configure Trigger Alarm | M |
| Enabled | Trigger Alarm Enabled | P |
| Period | Trigger Alarm Period | P |
| RepeatCount | Trigger Alarm Repeat Count | P |
| TimeSeconds | Trigger Alarm Time Seconds | P |
| TimeFraction | Trigger Alarm Time Fraction | P |
| Sources |  |  |
| Add | Add Trigger Source | M |
| Remove | Remove Trigger Source | M |
| RemoveAllCustomTriggerSources | Remove All Custom Trigger Sources | M |
| Count | Trigger Source Count | P |
| Name | Trigger Source Name | P |
| Item |  |  |
| Configure | Configure Trigger Source | M |
| Delay | Trigger Source Delay | P |
| Detection | Trigger Source Detection | P |
| Event ID | Trigger Source Event ID | P |
| Filter | Trigger Source Filter | P |
| Events |  |  |
| Add | Add Event | M |
| DisableAll | Disable All Events | M |
| Remove | Remove Event | M |
| RemoveAllCustomEvents | Remove All Custom Events | M |
| Count | Event Count | P |
| Name | Event Name | P |
| WiredOrBiasMode | Wired OR Bias Mode | P |
| Item |  |  |
| Configure | Configure Event | M |
| DestinationPath | Event Destination Path | P |
| Drive Mode | Event Drive Mode | P |
| Source | Event Source | P |
| Slope | Event Slope | P |
| EventLog |  |  |
| Clear | Clear Event Log | M |
| Enabled | Event Log Enabled | P |
| EntryCount | Event Log Entry Count | P |
| GetNextEntry | Get Next Event Log Entry | M |
| Time |  |  |
| GetSystemTime | Get System Time | M |
| IsMaster | Is Master | P |
| IsSynchronized | Is Synchronized | P |

### IviLxiSync COM Interfaces

IviLxiSync-interfaces contain interface reference properties for accessing the following IviLxiSync interfaces:

- IIviLixSyncArm

- IIviLixSyncEventLog

- IIviLixSyncEvents

- IIviLixSyncTime

- IIviLixSyncTrigger

The IIviLxiSyncArm interface contains interface reference properties for accessing additional the following IviLxiSync interfaces:

- IIviLxiSyncArmAlarms

- IIviLxiSyncArmSources

The IIviLxiSyncArmAlarms interface contains methods and properties for accessing a collection of objects that implement the IIviLxiSyncArmAlarm interface.

The IIviLxiSyncArmSources interface contains methods and properties for accessing a collection of objects that implement the IIviLxiSyncArmSource interface.

The IIviLxiSyncTrigger interface contains interface reference properties for accessing additional the following IviLxiSync interfaces:

- IIviLxiSyncTriggerAlarms

- IIviLxiSyncTriggerSources

The IIviLxiSyncTriggerAlarms interface contains methods and properties for accessing a collection of objects that implement the IIviLxiSyncTriggerAlarm interface.

The IIviLxiSyncTriggerSources interface contains methods and properties for accessing a collection of objects that implement the IIviLxiSyncTriggerSource interface.

The IIviLxiSyncEvents interface contains methods and properties for accessing a collection of objects that implement the IIviLxiSyncEvent interface.

### COM Interfaces

Table 12-4. Interface GUIDs lists the interfaces that this specification defines and their GUIDs.

| Table 12-4. Interface GUIDs |  |
| --- | --- |
| Interface | GUID |
| IIviLxiSync | 47ed5346-a398-11d4-ba58-000064657374 |
| IIviLxiSyncArm | 47ed5347-a398-11d4-ba58-000064657374 |
| IIviLxiSyncArmAlarm | 47ed5348-a398-11d4-ba58-000064657374 |
| IIviLxiSyncArmAlarms | 47ed5349-a398-11d4-ba58-000064657374 |
| IIviLxiSyncArmSource | 47ed534a-a398-11d4-ba58-000064657374 |
| IIviLxiSyncArmSources | 47ed534b-a398-11d4-ba58-000064657374 |
| IIviLxiSyncEvent | 47ed534c-a398-11d4-ba58-000064657374 |
| IIviLxiSyncEventLog | 47ed534d-a398-11d4-ba58-000064657374 |
| IIviLxiSyncEvents | 47ed534e-a398-11d4-ba58-000064657374 |
| IIviLxiSyncTime | 47ed534f-a398-11d4-ba58-000064657374 |
| IIviLxiSyncTrigger | 47ed5350-a398-11d4-ba58-000064657374 |
| IIviLxiSyncTriggerAlarm | 47ed5351-a398-11d4-ba58-000064657374 |
| IIviLxiSyncTriggerAlarms | 47ed5352-a398-11d4-ba58-000064657374 |
| IIviLxiSyncTriggerSource | 47ed5353-a398-11d4-ba58-000064657374 |
| IIviLxiSyncTriggerSources | 47ed5354-a398-11d4-ba58-000064657374 |

### COM Interface Reference Properties

Interface reference properties are used to navigate the IviLxiSync COM hierarchy. This section describes the interface reference properties that the IIviLxiSync, IIviLxiSyncArm, IIviLxiSyncTrigger, IIviLxiSyncEvents, and IIviLxiSyncEventLog.

| Table 12-5. Interface Reference Properties |  |
| --- | --- |
| DataType | Access |
| IIviLxiSyncArm | Arm |
| IIviLxiSyncArmAlarm | Alarm.Item() |
| IIviLxiSyncArmAlarms | Alarms |
| IIviLxiSyncArmSource | Sources.Item() |
| IIviLxiSyncArmSources | Sources |
| IIviLxiSyncEvent | Events.Item() |
| IIviLxiSyncEventLog | EventLog |
| IIviLxiSyncEvents | Events |
| IIviLxiSyncTime | Time |
| IIviLxiSyncTrigger | Trigger |
| IIviLxiSyncTriggerAlarm | Alarms.Item() |
| IIviLxiSyncTriggerAlarms | Alarms |
| IIviLxiSyncTriggerSource | Sources.Item() |
| IIviLxiSyncTriggerSources | Sources |

### COM Category

The IviLxiSync COM Category shall be “IviLxiSync”, and the Category ID (CATID) shall be {47ed515c-a398-11d4-ba58-000064657374}.

### COM Interface Accessibility

When an IviLxiSync-compliant IVI-COM driver is instantiated, a reference to the main driver class is returned.  A call to QueryInterface on this main class shall succeed for all IviLxiSync interfaces, except for interfaces that implement repeated capabilities as collections.

See IVI-3.1: Driver Architecture Specification for details on the relationship between IVI-COM driver classes, interfaces, and the QueryInterface function.

## C Function Hierarchy

The C function hierarchy is shown in the following table.

| Name or Class | Function Name |
| --- | --- |
| Configuration… |  |
| Arm… |  |
| Alarm… |  |
| Add Arm Alarm | IviLxiSync_AddArmAlarm |
| Configure Arm Alarm | IviLxiSync_ConfigureArmAlarm |
| Disable All Arm Alarms | IviLxiSync_DisableAllArmAlarms |
| Get Arm Alarm Name | IviLxiSync_GetArmAlarmName |
| Remove All Custom Arm Alarms | IviLxiSync_RemoveAllCustomArmAlarms |
| Remove Arm Alarm | IviLxiSync_RemoveArmAlarm |
| Source… |  |
| Add Arm Source | IviLxiSync_AddArmSource |
| Configure Arm Source | IviLxiSync_ConfigureArmSource |
| Disable All Arm Sources | IviLxiSync_DisableAllArmSources |
| Get Arm Source Name | IviLxiSync_GetArmSourceName |
| Remove All Custom Arm Sources | IviLxiSync_RemoveAllCustomArmSources |
| Remove Arm Source | IviLxiSync_RemoveArmSource |
| Trigger… |  |
| Alarm… |  |
| Add Trigger Alarm | IviLxiSync_AddTriggerAlarm |
| Configure Trigger Alarm | IviLxiSync_ConfigureTriggerAlarm |
| Disable All Trigger Alarms | IviLxiSync_DisableAllTriggerAlarms |
| Get Trigger Alarm Name | IviLxiSync_GetTriggerAlarmName |
| Remove All Trigger Alarms | IviLxiSync_RemoveAllTriggerAlarms |
| Source… |  |
| Add Trigger Source | IviLxiSync_AddTriggerSource |
| Configure Trigger Source | IviLxiSync_ConfigureTriggerSource |
| Get Trigger Source Name | IviLxiSync_GetTriggerSourceName |
| Remove All Custom Trigger Sources | IviLxiSync_RemoveAllCustomTriggerSources |
| Remove Trigger Source | IviLxiSync_RemoveTriggerSource |
| Event… |  |
| Add Event | IviLxiSync_AddEvent |
| Configure Event | IviLxiSync_ConfigureEvent |
| Disable All Events | IviLxiSync_DisableAllEvents |
| Get Event Name | IviLxiSync_GetEventName |
| Remove All Custom Events | IviLxiSync_RemoveAllCustomEvents |
| Remove Event | IviLxiSync_RemoveEvent |
| Event Log… |  |
| Clear Event Log | IviLxiSync_ClearEventLog |
| Get Next Event Log Entry | IviLxiSync_GetNextEventLogEntry |
| Time… |  |
| Get System Time | IviLxiSync_GetSystemTime |

## C Attribute Hierarchy

The IviLxiSync attribute hierarchy is shown in the following table.

|  | Table 12-7. C Attributes Hierarchy |  |
| --- | --- | --- |
|  | Category or Generic Attribute Name | C Defined Constant |
|  | Arm |  |
|  | Arm Count | IVILXISYNC_ATTR_ARM_COUNT |
|  | Arm Delay | IVILXISYNC_ATTR_ARM_DELAY |
|  | Alarm |  |
|  | Arm Alarm Count | IVILXISYNC_ATTR_ARM_ALARM_COUNT |
|  | Arm Alarm Enabled | IVILXISYNC_ATTR_ARM_ALARM_ENABLED |
|  | Arm Alarm Period | IVILXISYNC_ATTR_ARM_ALARM_PERIOD |
|  | Arm Alarm Repeat Count | IVILXISYNC_ATTR_ARM_ALARM_REPEAT_COUNT |
|  | Arm Alarm Time Seconds | IVILXISYNC_ATTR_ARM_ALARM_TIME_SECONDS |
|  | Arm Alarm Time Fraction | IVILXISYNC_ATTR_ARM_ALARM_TIME_FRACTION |
|  | Source |  |
|  | Arm Source Count | IVILXISYNC_ATTR_ARM_SOURCE_COUNT |
|  | Arm Source Detection | IVILXISYNC_ATTR_ARM_SOURCE_DETECTION |
|  | Arm Source Enabled | IVILXISYNC_ATTR_ARM_SOURCE_ENABLED |
|  | Arm Source Or Enabled | IVILXISYNC_ATTR_ARM_SOURCE_OR_ENABLED |
|  | Arm Source Event ID | IVILXISYNC_ATTR_ARM_SOURCE_EVENTID |
|  | Arm Source Filter | IVILXISYNC_ATTR_ARM_SOURCE_FILTER |
|  | Trigger |  |
|  | Trigger Count | IVILXISYNC_ATTR_TRIGGER_COUNT |
|  | Trigger Source | IVILXISYNC_ATTR_TRIGGER_SOURCE |
|  | Alarm |  |
|  | Trigger Alarm Count | IVILXISYNC_ATTR_TRIGGER_ALARM_COUNT |
|  | Trigger Alarm Enabled | IVILXISYNC_ATTR_TRIGGER_ALARM_ENABLED |
|  | Trigger Alarm Period | IVILXISYNC_ATTR_TRIGGER_ALARM_PERIOD |
|  | Trigger Alarm Repeat Count | IVILXISYNC_ATTR_TRIGGER_ALARM_REPEAT_COUNT |
|  | Trigger Alarm Time Seconds | IVILXISYNC_ATTR_TRIGGER_ALARM_TIME_SECONDS |
|  | Trigger Alarm Time Fraction | IVILXISYNC_ATTR_TRIGGER_ALARM_TIME_FRACTION |
|  | Source |  |
|  | Trigger Source Count | IVILXISYNC_ATTR_TRIGGER_SOURCE_COUNT |
|  | Trigger Source Delay | IVILXISYNC_ATTR_TRIGGER_SOURCE_DELAY |
|  | Trigger Source Detection | IVILXISYNC_ATTR_TRIGGER_SOURCE_DETECTION |
|  | Trigger Source Event ID | IVILXISYNC_ATTR_TRIGGER_SOURCE_EVENTID |
|  | Trigger Source Filter | IVILXISYNC_ATTR_TRIGGER_SOURCE_FILTER |
|  | Event |  |
|  | Event Count | IVILXISYNC_ATTR_EVENT_COUNT |
|  | Event Wired OR Bias Mode | IVILXISYNC_ATTR_EVENT_WIRED_OR_BIAS_MODE |
|  | Event Destination Path | IVILXISYNC_ATTR_EVENT_DESTINATION_PATH |
|  | Event Drive Mode | IVILXISYNC_ATTR_EVENT_DRIVE_MODE |
|  | Event Slope | IVILXISYNC_ATTR_EVENT_SLOPE |
|  | Event Source | IVILXISYNC_ATTR_EVENT_SOURCE |
|  | Event Log |  |
|  | Event Log Enabled | IVILXISYNC_ATTR_EVENT_LOG_ENABLED |
|  | Event Log Entry Count | IVILXISYNC_ATTR_EVENT_LOG_ENTRY_COUNT |
|  | Time |  |
|  | Is Time Master | IVILXISYNC_ATTR_IS_TIME_MASTER |
|  | Is Time Synchronized | IVILXISYNC_ATTR_IS_TIME_SYNCHRONIZED |