IVI-4.4: IviDCPwr Class Specification September 24, 2015 Edition Revision 3.0

- Important Information

The IviDCPwr Class Specification (IVI-4.4) is authored by the IVI Foundation member companies. For a vendor membership roster list, please visit the IVI Foundation web site at www.ivifoundation.org.

The IVI Foundation wants to receive your comments on this specification. You can contact the Foundation through the web site at www.ivifoundation.org.

- Warranty

The IVI Foundation and its member companies make no warranty of any kind with regard to this material, including, but not limited to, the implied warranties of merchantability and fitness for a particular purpose. The IVI Foundation and its member companies shall not be liable for errors contained herein or for incidental or consequential damages in connection with the furnishing, performance, or use of this material.

- Trademarks

Product and company names listed are trademarks or trade names of their respective companies.

No investigation has been made of common-law trademark rights in any work.

1	Overview of the IviDCPwr Specification	8

1.1	Introduction	8

1.2	IviDCPwr Class Overview	8

1.3	References	8

1.4	Definitions of Terms and Acronyms	9

2	IviDCPwr Class Capabilities	10

2.1	Introduction	10

2.2	IviDCPwr Group Names	10

2.3	Repeated Capability Names	10

2.3.1	OutputChannel	10

2.4	Boolean Attribute and Parameter Values	11

2.5	.NET Namespace	11

2.6	.NET IviDCPwr Session Factory	12

3	General Requirements	14

3.1	Minimum Class Compliance	14

3.1.1	Disable	14

3.2	Capability Group Compliance	14

4	IviDCPwrBase Capability Group	15

4.1	Overview	15

4.2	IviDCPwrBase Attributes	17

4.2.1	Current Limit	18

4.2.2	Current Limit Behavior	19

4.2.3	Output Enabled	21

4.2.4	OVP Enabled	22

4.2.5	OVP Limit	23

4.2.6	Voltage Level	24

4.2.7	Output Channel Count	25

4.2.8	Output Channel Item (IVI-COM and IVI.NET only)	26

4.2.9	Output Channel Name (IVI-COM and IVI.NET only)	27

4.3	IviDCPwrBase Functions	28

4.3.1	Configure Current Limit	29

4.3.2	Configure Output Enabled (IVI-C only)	30

4.3.3	Configure Output Range	31

4.3.4	Configure OVP	33

4.3.5	Configure Voltage Level (IVI-C only)	34

4.3.6	Get Output Channel Name (IVI-C only)	35

4.3.7	Query Current Limit Max	36

4.3.8	Query Voltage Level Max	37

4.3.9	Query Output State	38

4.3.10	Reset Output Protection	40

4.4	IviDCPwrBase Behavior Model	41

5	IviDCPwrTrigger Extension Group	42

5.1	IviDCPwrTrigger Overview	42

5.2	IviDCPwrTrigger Attributes	42

5.2.1	Trigger Source	43

5.2.2	Triggered Current Limit	47

5.2.3	Triggered Voltage Level	48

5.3	IviDCPwrTrigger Functions	49

5.3.1	Abort	50

5.3.2	Configure Triggered Current Limit (IVI-C only)	51

5.3.3	Configure Triggered Voltage Level (IVI-C only)	52

5.3.4	Configure Trigger Source (IVI-C only)	53

5.3.5	Initiate	54

5.4	IviDCPwrTrigger Behavior Model	55

6	IviDCPwrSoftwareTrigger Extension Group	56

6.1	IviDCPwrSoftwareTrigger Overview	56

6.2	IviDCPwrTrigger Functions	56

6.2.1	Send Software Trigger	56

6.3	IviDCPwrSoftwareTrigger Behavior Model	56

6.4	IviDCPwrSoftwareTrigger Compliance Notes	56

7	IviDCPwrMeasurement Extension Group	57

7.1	IviDCPwrMeasurement Overview	57

7.2	IviDCPwrMeasurement Functions	57

7.2.1	Measure	58

7.3	IviDCPwrMeasurement Behavior Model	60

8	IVIDCPwr Attribute ID Definitions	61

9	IVIDCPwr Attribute Value Definitions	62

10	IviDCPwr Function Parameter Value Definitions	65

11	Error and Completion Code Value Definitions	67

12	IviDCPwr Hierarchies	68

12.1	IviDCPwr .NET Hierarchy	68

12.1.1	IviDCPwr .NET Interface Reference Properties	68

12.2	IviDCPwr COM Hierarchy	69

12.2.1	IviDCPwr COM Interfaces	69

12.2.2	IviDCPwr COM Interface Reference Properties	70

12.2.3	IviDCPwr COM Category	70

12.3	IviDCPwr C Function Hierarchy	71

12.4	IviDCPwr C Attribute Hierarchy	72

Appendix A	Specific Driver Development Guidelines	73

A.1	Introduction	73

A.2	Disabling Unused Extensions	73

A.3	Special Considerations for Configure Output Range	74

A.4	Special Considerations for Initiate	74

A.5	Special Considerations for Current Limit and Current Limit Behavior	75

Appendix B.	Interchangeability Checking Rules	76

B.1	Introduction	76

B.2	When to Perform Interchangeability Checking	76

B.3	Interchangeability Checking Rules	76

IviDCPwr Class Specification

IviDCPwr Revision History

This section is an overview of the revision history of the IviDCPwr specification.

| Table 1. IviDCPwr Class Specification Revisions |  |  |
| --- | --- | --- |
| Revision Number | Date of Revision | Revision Notes |
| Revision 1.0 draft | October 1, 1999 | First approved version. Original draft. |
| Revision 2.0a draft | April 31, 2001 | First draft to include COM requirements |
| Revision 2.1vc1 draft | July 30, 2001 | Voting candidate 1 review draft. |
| Revision 2.1vc2 draft | August 30, 2001 | Voting candidate 2 review draft. |
| Revision 2.1vc3 draft | November, 2001 | Voting candidate 3 review draft. |
| Revision 2.1vc5 draft | December, 2001 | Voting candidate 5 review draft. Changes based on discussion at December IVI meeting. |
| Revision 2.0vc6 draft | January, 2002 | Voting candidate 6 review draft. Changes based on comments from NI. |
| Revision 2.0vc7 draft | March, 2002 | Added latest IDL. |
| Revision 2.0vc7 draft | March, 2002 | IVI_CLASS_PUBLIC_ATTR_BASE changed to IVI_CLASS_ATTR_BASE. |
| Revision 2.0 Final | March, 2002 | Approved Specification |
| Revision 2.0 Final | March, 2008 | Editorial change to update the IVI Foundation contact information in the Important Information section to remove obsolete address information and refer only to the IVI Foundation web site. |
| Revision 2.0 Final | April 2009 | Editorial change to update repeated capabilities section to include both qualified and unqualified repeated capability names. |
| Revision 3.0 | June 9, 2010 | Incorporated IVI.NET |
| Revision 3.0 | August 25, 2011 | Editorial IVI.NET change. Change references to process-wide locking to AppDomain-wide locking. Add an overload to the Create factory method that takes locking related parameters. |
| Revision 3.0 | September 24, 2015 | Editorial Change – Clarified the use of one-based index and COM, and zero-based index for .NET for repeated capabilities in section 4.2.9. |

API Versions

| Architecture | Drivers that comply with version 3.0 comply with all of the versions below |
| --- | --- |
| C | 2.0, 3.0 |
| COM | 2.0, 3.0 |
| .NET | 3.0 |

Drivers that comply with this version of the specification also comply with earlier, compatible, versions of the specification as shown in the table above.  The driver may benefit by advertising that it supports all the API versions listed in the table above.

# Overview of the IviDCPwr Specification

## Introduction

This specification defines the IVI class for DC power supplies. The IviDCPwr class is designed to support the typical DC power supply as well as common extended functionality found in more complex instruments. This section summarizes the IviDCPwr Specification itself and contains general information that the reader may need in order to understand, interpret, and implement aspects of this specification. These aspects include the following:

- IviDCPwr Class Overview

- The definitions of terms and acronyms

- References

## IviDCPwr Class Overview

This specification describes the IVI class for DC power supplies. The IviDCPwr class is designed to support the typical DC Power supply as well as common extended functionality found in more complex instruments. The IviDCPwr class conceptualizes a DC power supply as an instrument capable of generating a DC power signal.

The IviDCPwrBase capability group provides the capability to configure a power supply for basic signal output. This includes setting the output range, the output voltage level, the over-voltage protection level, the current limit, enabling or disabling over-voltage protection, setting the current limit behavior, and enabling or disabling outputs. The IviDCPwrBase capability group is described in section 4, IviDCPwrBase Capability Group.

The IviDCPwrTrigger extension provides the capability to make changes to the output signal based on a trigger event. The driver can configure the trigger source, the triggered voltage level, and the triggered current limit. This extension group is described in section 5, IviDCPwrTrigger Extension Group.

The IviDCPwrSoftwareTrigger extension provides the capability to make changes to the output signal based on a software trigger event. Sending a software trigger to the instrument causes it to change the output signal. This extension group is described in section 6, IviDCPwrSoftwareTrigger Extension Group.

The IviDCPwrMeasurement extension provides the capability to take measurements on the output signal. With this extension, the driver can take immediate measurements of values like the DC Voltage and DC Current. This extension group is described in section 7, IviDCPwrMeasurement Extension Group.

## References

Several other documents and specifications are related to this specification. These other related documents are the following:

- IVI Charter Document

- IVI-3.1 - Driver Architecture Specification

- IVI-3.2 - Inherent Capabilities Specification

- IVI-3.3 – Standard Cross Class Capabilities

- IVI-3.18: IVI.NET Utility Classes and Interfaces Specification

- VPP-3.x - VXIplug&play Instrument Driver Specifications

- VPP-4.x - Virtual Instrument Software Architecture (VISA) Specifications

## Definitions of Terms and Acronyms

Refer to IVI-5: Glossary for a description of the terms and acronyms used in this specification. This specification does not define any additional terms.

# IviDCPwr Class Capabilities

## Introduction

The IviDCPwr specification divides DC power supply capabilities into a base capability group and multiple extension capability groups. Each capability group is discussed in a separate section. This section defines names for each capability group and gives an overview of the information presented for each capability group.

## IviDCPwr Group Names

The capability group names for the IviDCPwr class are defined in the following table. The group name is used to represent a particular capability group and is returned as one of the possible group names from the Group Capabilities attribute.

| Table 2-1. IviDCPwr Group Names |  |
| --- | --- |
| Group Name | Description |
| IviDCPwrBase | Base capabilities of the IviDCPwr specification. This group supports the ability to generate a DC power signal, to specify output limits, and to control the behavior of the power supply when the output is greater than or equal to one of the limits. |
| IviDCPwrTrigger | This group supports the ability to make changes to the output signal based on a trigger event. |
| IviDCPwrSoftwareTrigger | This group supports the ability to make changes to the output signal based on a software trigger event. |
| IviDCPwrMeasurement | This group supports the ability to query the instrument for the measurement characteristics of the output signal. |

Refer to Section 16, Class Specification Layout, in IVI-3.4: API Style Guide for a description of the Capability Group Section Layout.

## Repeated Capability Names

The IviDCPwr Class Specification defines one repeated capability.  Refer to the sections of IVI-3.1, Driver Architecture Specification that deal with repeated capabilities.  The relevant sections are Section 2.7, Repeated Capabilities, Section 4.1.9, Repeated Capabilities, Section 4.2.5, Repeated Capabilities, Section 4.3.9, Repeated Capabilities, and Section 5.8, Repeated Capability Identifiers and Selectors.

- OutputChannel

### OutputChannel

In the configuration store, the name for the channel repeated capability shall be exactly one of “OutputChannel” or “IviDCPwrOutputChannel”.  Drivers that implement multiple repeated capabilities with the name “OutputChannel” shall use the latter form to disambiguate the names.

## Boolean Attribute and Parameter Values

This specification uses True and False as the values for Boolean attributes and parameters. The following table defines the identifiers that are used for True and False in the IVI.NET, IVI-COM, and IVI-C architectures.

| Boolean Value | IVI.NET Identifier | IVI-COM Identifier | IVI-C Identifier |
| --- | --- | --- | --- |
| True | true | VARIANT_TRUE | VI_TRUE |
| False | false | VARIANT_FALSE | VI_FALSE |

## .NET Namespace

The .NET namespace for the IviDCPwr class is Ivi.DCPwr.

## .NET IviDCPwr Session Factory

The IviDCPwr .NET assembly contains a factory method called Create for creating instances of IviDCPwr class-compliant IVI.NET drivers from driver sessions and logical names.  Create is a static method accessible from the static IviDCPwr class.

Refer to IVI-3.5: Configuration Server Specification for a description of how logical names and session names are defined in the configuration store.

Refer to Section 8, IVI.NET Specific Driver Constructor, of IVI-3.2: Inherent Capabilities Specification, for more details on how the idQuery, reset, and options parameters affect the instantiation of the driver.

Refer to Section 4.3.11, Multithread Safety, of IVI-3.1: Driver Architecture Specification for a complete description of IVI.NET driver locking.  Refer to Section 8, Table 8.2 Required Lock Type Behavior for Drivers With the Same Access Key, of  IVI-3.2, Inherent Capability Specification, for an explanation of how the values for lockType and accessKey are used to determine the kind of multithreaded lock to use for the driver instance.

.NET Method Prototype

IIviDCPwr Ivi.DCPwr.Create(String name);

IIviDCPwr Ivi.DCPwr.Create(String name,

Boolean idQuery,

Boolean reset);

IIviDCPwr Ivi.DCPwr.Create(String name,

Boolean idQuery,

Boolean reset,

String options);

IIviDCPwr Ivi.DCPwr.Create(String resourceName,

Boolean idQuery,

Boolean reset,

LockType lockType,

String accessKey,

String options);

Parameters

| Inputs | Description | Base Type |  |
| --- | --- | --- | --- |
| name | A session name or a logical name that points to a session that uses an IVI.NET IviDCPwr class-compliant driver. | String |  |
| idQuery | Specifies whether to verify the ID of the instrument. The default is False. | Boolean |  |
| reset | Specifies whether to reset the instrument. The default is False. | Boolean |  |
| lockType | Specifies whether to use AppDomain-wide locking or machine-wide locking. |  | Ivi.Driver.LockType |
| accessKey | Specifies a user-selectable access key to identify the lock. Driver instances that are created with the same accessKey will be protected from simultaneous access by multiple threads within an AppDomain or across AppDomains, depending upon the value of the lockType parameter. |  | String |
| options | A string that allows the user to specify the initial values of certain inherent attributes. The default is an empty string. | String |  |
| Outputs | Description | Base Type |  |
| Return Value | Interface pointer to the IIviDCPwr interface of the driver referenced by session. | IIviDCPwr |  |

      - Defined Values

| Name | Description |  |  |
| --- | --- | --- | --- |
|  |  | Language | Identifier |
| AppDomain | The lock is AppDomain-wide. |  |  |
|  |  | .NET | Ivi.Driver.LockType.AppDomain |
| Machine | The lock is machine-wide. |  |  |
|  |  | .NET | Ivi.Driver.LockType.Machine |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

Usage

To create a driver that implements the IviDCPwr instrument class API from the logical name “My LogicalName” use the following code:

IIviDCPwr dcPwr = IviDCPwr.Create(“MyLogicalName”);

In this case, the ID of the instrument will not be verified, the instrument will not be reset, and options will be supplied from the configuration store and/or driver defaults.

# General Requirements

This section describes the general requirements a specific driver shall meet in order to be compliant with this specification. In addition, it provides general requirements that specific drivers shall meet in order to comply with a capability group, attribute, or function.

## Minimum Class Compliance

To be compliant with the IviDCPwr Class Specification, a specific driver shall implement the inherent capabilities that IVI-3.2: Inherent IVI Capabilities Specification defines, and the IviDCPwrBase capability group.

### Disable

Refer to IVI-3.2: Inherent Capabilities Specification for the prototype of this function.

The Disable function shall cause the DC Power Supply to apply the minimum amount of power possible at the output terminals. Setting the voltage to a value close to zero, setting the current limit to value close to zero, or physically disconnecting the power supply from the output terminals meets this requirement. Other techniques are also allowed.

## Capability Group Compliance

IVI-3.1: Driver Architecture Specification defines the general rules for a specific driver to be compliant with a capability group.

# IviDCPwrBase Capability Group

## Overview

The IviDCPwrBase capability group supports the most basic DC power supply capabilities. The user can enable or disable outputs, specify the DC voltage to generate, specify output limits, and control the behavior of the power supply when the output is greater than or equal to one of the limits.

This specification uses the following terms to describe the power supply’s output: Voltage Level, OVP Limit, Current Limit, Current Limit Behavior, Constant Voltage Mode, Constant Current Mode, and Unregulated Mode.

Voltage Level – The DC voltage the power supply attempts to generate. The user configures the voltage level with the Voltage Level attribute.

OVP – OVP is an acronym for Over-Voltage Protection.

OVP Limit and OVP Enabled – If the OVP limit is enabled, the power supply disables the output when the output voltage is greater than or equal to the OVP limit. The user configures the OVP limit with the OVP Limit attribute, and enables or disables the OVP limit with the OVP Enabled attribute.

Current Limit and Current Limit Behavior – The current limit behavior determines the behavior of the instrument when the output current is greater than or equal to the current limit. When the current limit behavior is trip, the power supply disables the output when the output current is greater than or equal to the current limit. When the current limit behavior is regulate, the power supply restricts the output voltage such that the output current is not greater than the current limit. The user configures the current limit and current limit behavior with the Current Limit and Current Limit Behavior attributes.

Constant Voltage Mode – The power supply is said to be in the constant voltage mode when the power supply’s output signal reaches the voltage level before it reaches the current limit. In the constant voltage mode, the power supply’s output voltage remains constant at the voltage level and its output current can vary.

Constant Current Mode – The power supply is said to be in the constant current mode when the power supply’s output signal reaches the current limit before it reaches the voltage level, and the current limit behavior is set to regulate. In the constant current mode, the power supply’s output current remains constant at the current limit and its output voltage varies.

Unregulated Mode – The power supply is said to be in the unregulated mode when the power supply’s output signal reaches neither the voltage level or the current limit. In the unregulated mode, the power supply’s output current and output voltage vary.

The signal that the power supply produces depends on the values of the voltage level, OVP limit, and current limit that the user supplies, and the impedance of the load to which the power supply is attached. Therefore, the power supply might not produce the exact voltage or current that the user configures. The following diagram shows the possible output scenarios.

- Figure 4-1. Power Supply Output Diagram

At Point 1, the power supply’s output has reached the voltage level before it reached the current limit. This is an example of the power supply operating in the constant voltage mode. Note that any point on the vertical line defined by the voltage level would also cause the power supply to operate in the constant voltage mode.

At Point 2, the power supply’s output has reached the current limit before it reached the voltage level. If the current limit behavior is set to regulate, this is an example of the power supply operating in the constant current mode. If the current limit behavior is set to trip, the power supply disables the output. Note that any point on the horizontal line defined by the current limit would also cause the power supply to operate in the constant current mode when the current limit behavior is set to regulate.

At point 3, the power supply’s output has reached neither the voltage level or the current limit. This is an example of the power supply operating in the unregulated mode. Note that any point within the rectangle defined by the voltage level and current limit would also cause the power supply to operate in the unregulated mode.

At point 4, the power supply’s output has reached the OVP limit. If OVP is enabled, the power supply disables the output. Note that any point on the vertical line defined by the OVP limit would also cause the power supply to disable the output when OVP is enabled.

## IviDCPwrBase Attributes

The IviDCPwrBase capability group defines the following attributes:

- Current Limit

- Current Limit Behavior

- Output Enabled

- OVP Enabled

- OVP Limit

- Voltage Level

- Output Channel Count

- Output Channel Item (IVI-COM and IVI.NET only)

- Output Channel Name (IVI-COM and IVI.NET only)

This section describes the behavior and requirements of each attribute. The actual value for each attribute ID is defined in section 8, IVIDCPwr Attribute ID Definitions.

### Current Limit

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 | R/W | OutputChannel | None | Configure Current Limit |

.NET Property Name

Outputs[].CurrentLimit

COM Property Name

Outputs.Item().CurrentLimit

C Constant Name

IVIDCPWR_ATTR_CURRENT_LIMIT

Description

Specifies the output current limit. The units are Amps.

The value of the Current Limit Behavior attribute determines the behavior of the power supply when the output current is equal to or greater than the value of this attribute.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

### Current Limit Behavior

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | R/W | OutputChannel | None | Configure Current Limit |

.NET Property Name

Outputs[].CurrentLimitBehavior

.NET Enumeration Name

Ivi.DCPwr.CurrentLimitBehavior

COM Property Name

Outputs.Item().CurrentLimitBehavior

COM Enumeration Name

IviDCPwrCurrentLimitEnum

C Constant Name

IVIDCPWR_ATTR_CURRENT_LIMIT_BEHAVIOR

Description

Specifies the behavior of the power supply when the output current is equal to or greater than the value of the Current Limit attribute.

Defined Values

| Name | Description |  |  |  |
| --- | --- | --- | --- | --- |
|  |  | Language |  | Identifier |
| Current Trip | The power supply disables the output when the output current is equal to or greater than the value of the Current Limit attribute. |  |  |  |
|  |  | C | IVIDCPWR_VAL_CURRENT_TRIP |  |
|  |  | COM | IviDcPwrCurrentLimitTrip |  |
|  |  | .NET | Ivi.DCPwr.CurrentLimitBehavior.Trip |  |
| Current Regulate | The power supply restricts the output voltage such that the output current is not greater than the value of the Current Limit attribute. |  |  |  |
|  |  | C | IVIDCPWR_VAL_CURRENT_REGULATE |  |
|  |  | COM | IviDcPwrCurrentLimitRegulate |  |
|  |  | .NET | Ivi.DCPwr.CurrentLimitBehavior.Regulate |  |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

Compliance Notes

- If an IVI-C specific driver defines additional values for this attribute, the actual values must be greater than or equal to IVIDCPWR_VAL_CURRENT_LIMIT_BEHAVIOR_SPECIFIC_EXT_BASE.

- If an IVI-C class driver defines additional values for this attribute, the actual values shall be greater than or equal to IVIDCPWR_VAL_CURRENT_LIMIT_BEHAVIOR_CLASS_EXT_BASE and less than IVIDCPWR_VAL_CURRENT_LIMIT_BEHAVIOR_SPECIFIC_EXT_BASE.

- When an IVI-COM specific driver implements this attribute with additional elements in its instrument specific interfaces, the actual values of the additional elements shall be greater than or equal to Current Limit Behavior Specific Ext Base.

### Output Enabled

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViBoolean | R/W | OutputChannel | None | Configure Output Enabled |

.NET Property Name

Outputs[].Enabled

COM Property Name

Outputs.Item().Enabled

C Constant Name

IVIDCPWR_ATTR_OUTPUT_ENABLED

Description

If true, the signal the power supply produces appears at the output connector. If false, the signal the power supply produces does not appear at the output connector.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

Compliance Notes

Instrument drivers shall support the value True on all outputs.

### OVP Enabled

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViBoolean | R/W | OutputChannel | None | Configure OVP |

.NET Property Name

Outputs[].OvpEnabled

COM Property Name

Outputs.Item().OVPEnabled

C Constant Name

IVIDCPWR_ATTR_OVP_ENABLED

Description

Specifies whether the power supply provides over-voltage protection. If this attribute is set to True, the power supply disables the output when the output voltage is greater than or equal to the value of the OVP Limit attribute.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

### OVP Limit

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 | R/W | OutputChannel | None | Configure OVP |

.NET Property Name

Outputs[].OvpLimit

COM Property Name

Outputs.Item().OVPLimit

C Constant Name

IVIDCPWR_ATTR_OVP_LIMIT

Description

Specifies the voltage the power supply allows. The units are Volts.

If the OVP Enabled attribute is set to True, the power supply disables the output when the output voltage is greater than or equal to the value of this attribute.

If the OVP Enabled is set to False, this attribute does not affect the behavior of the instrument.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

### Voltage Level

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 | R/W | OutputChannel | None | Configure Voltage Level |

.NET Property Name

Outputs[].VoltageLevel

COM Property Name

Outputs.Item().VoltageLevel

C Constant Name

IVIDCPWR_ATTR_VOLTAGE_LEVEL

Description

Specifies the voltage level the DC power supply attempts to generate. The units are Volts.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

### Output Channel Count

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 | RO | IviDCPwrOutput | None | None |

.NET Property Name

Outputs.Count

This property is inherited from IIviRepeatedCapabilityCollection.

COM Property Name

Outputs.Count

C Constant Name

IVIDCPWR_ATTR_CHANNEL_COUNT

Description

Returns the number of currently available outputs.  The count returned includes any of the supported reserved repeated capability names defined in Section 2.3, Repeated Capability Names as well as any custom repeated capability identifiers.  The C Constant Name uses CHANNEL while the COM Property Name is tied to the Outputs collection.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

### Output Channel Item (IVI-COM and IVI.NET only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| IIviDCPwrOutput* | RO | IviDCPwrOutput | None | N/A |

.NET Property Name

Outputs[String name]

This indexer is inherited from IIviRepeatedCapabilityCollection.  The name parameter uniquely identifies a particular output channel in the outputs collection.

COM Property Name

Outputs.Item([in] BSTR OutputName)

C Constant Name

N/A

Description

Output Channel Item uniquely identifies an output channel in the outputs collection. It returns an interface pointer which can be used to control the attributes and other functionality of that output channel.

The Item property takes an output name.  If the user passes an invalid value for the output name parameter, the property returns an error.

Valid names include physical repeated capability identifiers and virtual repeated capability identifiers.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

### Output Channel Name (IVI-COM and IVI.NET only)

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViString | RO | IviDCPwrOutput | None | N/A |

.NET Property Name

Outputs[].Name

This property is inherited from IIviRepeatedCapabilityIdentification.

COM Property Name

Outputs.Name ([in] LONG Index)

C Constant Name

N/A

(Use the Get Output Channel Name function.)

Description

This attribute returns the repeated capability identifier defined by specific driver for the output channel that corresponds to the index that the user specifies.  If the driver defines a qualified Output Channel name, this property returns the qualified name.

In COM, the index is one-based.  In .NET, the index is zero-based.

For C and COM, valid values for the Index parameter are between one and the value of the Output Channel Count attribute, inclusive. If the user passes an invalid value for the Index parameter, the value of this attribute is an empty string, and the property returns an error.

If the value that the user passes for the Index parameter is less than one or greater than the value of the Output Channel Count, the attribute returns an empty string for the value and returns an error.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

## IviDCPwrBase Functions

The IviDCPwrBase capability group defines the following functions:

- Configure Current Limit

- Configure Output Enabled (IVI-C only)

- Configure Output Range

- Configure OVP

- Configure Voltage Level (IVI-C only)

- Get Output Channel Name (IVI-C only)

- Query Current Limit Max

- Query Voltage Level Max

- Query Output State

- Reset Output Protection

This section describes the behavior and requirements of each function.

### Configure Current Limit

Description

This function configures the current limit. It specifies the output current limit value and the behavior of the power supply when the output current is greater than or equal to that value.

See the definition of the Current Limit Behavior attribute for defined values for the behavior parameter.

.NET Method Prototype

void Outputs[].ConfigureCurrentLimit(CurrentLimitBehavior behavior,

Double limit);

COM Method Prototype

HRESULT Outputs.Item().ConfigureCurrentLimit(
[in] IviDCPwrCurrentLimitBehaviorEnum Behavior,
[in] DOUBLE Limit);

C Prototype

ViStatus IviDCPwr_ConfigureCurrentLimit (ViSession Vi,
                                         ViConstString ChannelName,
                                         ViInt32 Behavior,
                                         ViReal64 Limit);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The name of the output on which to configure the current limit. | ViConstString |
| Behavior | Specifies the behavior of the power supply when the output current is greater than or equal to the value of the Limit parameter. The driver uses this value to set the Current Limit Behavior attribute. See the attribute description for more details. | ViInt32 |
| Limit | Specifies the power supply’s output current limit. The driver uses this value to set the Current Limit attribute. See the attribute description for more details. | ViReal64 |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

### Configure Output Enabled (IVI-C only)

Description

Configures whether the signal that the power supply produces appears at the output connector.

.NET Method Prototype

N/A

(use the Outputs[].Enabled property)

COM Method Prototype

N/A

(use the Outputs.Item().Enabled property)

C Prototype

ViStatus IviDCPwr_ConfigureOutputEnabled (ViSession Vi,
                                          ViConstString ChannelName,
                                          ViBoolean Enabled);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The name of the output to enable or disable. | ViConstString |
| Enabled | Specifies whether the signal the power supply produces appears at the output connector. The driver uses this value to set the Output Enabled attribute. See the attribute description for more details. | ViBoolean |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

### Configure Output Range

Description

Configures the power supply’s output range on an output.  One parameter specifies whether to configure the voltage or current range, and the other parameter is the value to which to set the range.

Setting a voltage range can invalidate a previously configured current range.  Setting a current range can invalidate a previously configured voltage range.

.NET Method Prototype

void Outputs[].ConfigureRange(Ivi.DCPwr.RangeType rangeType, Double range);

COM Method Prototype

HRESULT Outputs.Item().ConfigureRange([in] IviDCPwrRangeTypeEnum RangeType,
           [in] DOUBLE Range);

C Prototype

ViStatus IviDCPwr_ConfigureOutputRange (ViSession Vi,
                                        ViConstString ChannelName,
                                        ViInt32 RangeType,
                                        ViReal64 Range);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The name of the output on which to configure the output range. | ViConstString |
| RangeType | The kind of range to be configured. | ViInt32 |
| Range | The range in which the power supply operates. The instrument driver coerces this value to the closest value the instrument supports that is greater than or equal to the value specified. | ViReal64 |

Defined Values for RangeType Parameter

| Name | Description |  |  |  |
| --- | --- | --- | --- | --- |
|  |  | Language |  | Identifier |
| Voltage | Voltage range is set by the Range parameter. |  |  |  |
|  |  | .NET | RangeType.Voltage |  |
|  |  | C | IVIDCPWR_VAL_RANGE_VOLTAGE |  |
|  |  | COM | IviDCPwrRangeVoltage |  |
| Current | Current range is set by the Range parameter. |  |  |  |
|  |  | C | IVIDCPWR_VAL_RANGE_CURRENT |  |
|  |  | .NET | RangeType.Current |  |
|  |  | COM | IviDCPwrRangeCurrent |  |

Return Values (C/COM)

The IVI-3.2: Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

Compliance Notes

- If an IVI-C class driver defines additional values for the RangeType parameter, the actual values shall be greater than or equal to IVIDCPWR_VAL_RANGE_TYPE_CLASS_EXT_BASE and less than IVIDCPWR_VAL_RANGE_TYPE_SPECIFIC_EXT_BASE.

- If an IVI-C specific driver defines additional values for the RangeType parameter, the actual values shall be greater than or equal to IVIDCPWR_VAL_RANGE_TYPE_SPECIFIC_EXT_BASE.

### Configure OVP

Description

Configures the over-voltage protection.  It specifies the over-voltage limit and the behavior of the power supply when the output voltage is greater than or equal to that value.

When the Enabled parameter is False, the Limit parameter does not affect the instrument’s behavior, and the driver does not set the OVP Limit attribute.

.NET Method Prototype

void Outputs[].ConfigureOvp(Boolean enabled, Double limit);

COM Method Prototype

HRESULT Outputs.Item().ConfigureOVP([in] VARIANT_BOOL Enabled,
         [in] DOUBLE Limit);

C Prototype

ViStatus IviDCPwr_ConfigureOVP (ViSession Vi,
                                ViConstString ChannelName,
                                ViBoolean Enabled,
                                ViReal64 Limit);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The name of the output on which to configure over-voltage protection. | ViConstString |
| Enabled | Specifies the behavior of the power supply when the ouptut voltage is greater than or equal to the value of the Limit parameter. The driver uses this value to set the OVP Enabled attribute. See the attribute description for more details. | ViBoolean |
| Limit | Specifies the power supply’s over-voltage protection limit. The driver uses this value to set the OVP Limit attribute. See the attribute description for more details. | ViReal64 |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

### Configure Voltage Level (IVI-C only)

Description

Configures the voltage level the power supply attempts to generate.

.NET Method Prototype

N/A

(use the Outputs[].VoltageLevel property)

COM Method Prototype

N/A

(use the Outputs.Item().VoltageLevel property)

C Prototype

ViStatus IviDCPwr_ConfigureVoltageLevel (ViSession Vi,
                                         ViConstString ChannelName,
                                         ViReal64 Level);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The name of the output on which to configure the voltage level. | ViConstString |
| Level | Specifies the voltage level the power supply attempts to generate. The driver uses this value to set the Voltage Level attribute. See the attribute description for more details. | ViReal64 |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

### Get Output Channel Name (IVI-C only)

Description

This function returns the specific driver defined output name that corresponds to the one-based index that the user specifies.  If the driver defines a qualified Output Channel name, this property returns the qualified name.  If the value that the user passes for the Index parameter is less than one or greater than the value of the Output Channel Count attribute, the function returns an empty string in the Name parameter and returns an error.

The C prototype uses Channel in the function name rather than the name of the repeated capability.

.NET Method Prototype

N/A

(use the Outputs[].Name property)

COM Method Prototype

N/A

(use the Outputs.Name property)

C Prototype

ViStatus IviDCPwr_GetChannelName (ViSession Vi, 
      ViInt32 Index,
      ViInt32 NameBufferSize,
      ViChar Name[]);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| Index | A one-based index that defines which name to return. | ViInt32 |
| NameBufferSize | The number of bytes in the ViChar array that the user specifies for the Name parameter. | ViInt32 |

| Outputs | Description | Base Type |
| --- | --- | --- |
| Name | A user-allocated (for IVI-C) or driver-allocated (for IVI-COM) buffer into which the driver stores the output name. The caller may pass VI_NULL for this parameter if the NameBufferSize parameter is 0. | ViChar[] |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

### Query Current Limit Max

Description

This function returns the maximum programmable current limit that the power supply accepts for a particular voltage level on an output.

.NET Method Prototype

Double Outputs[].QueryCurrentLimitMax(Double voltageLevel);

COM Method Prototype

HRESULT Outputs.Item().QueryCurrentLimitMax(
[in] DOUBLE VoltageLevel,
[out, retval] DOUBLE *MaxCurrentLimit);

C Prototype

ViStatus IviDCPwr_QueryMaxCurrentLimit (ViSession  Vi,
                                        ViConstString ChannelName,
                                        ViReal64 VoltageLevel,
                                        ViReal64* MaxCurrentLimit);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The name of the output on which to query the maximum current limit. | ViConstString |
| VoltageLevel | The voltage level for which to determine the maximum programmable current limit. | ViReal64 |

| Outputs | Description | Base Type |
| --- | --- | --- |
| Return value (.NET) | Returns the maximum programmable current limit for the specified voltage level. | ViReal64 |
| MaxCurrentLimit (C/COM) | Returns the maximum programmable current limit for the specified voltage level. | ViReal64* |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

### Query Voltage Level Max

Description

This function returns the maximum programmable voltage level that the power supply accepts for a particular current limit on an output.

.NET Method Prototype

Double Outputs[].QueryVoltageLevelMax(Double currentLimit);

COM Method Prototype

HRESULT Outputs.Item().QueryVoltageLevelMax(
[in] DOUBLE CurrentLimit,
[out, retval] DOUBLE *MaxVoltageLevel);

C Prototype

ViStatus IviDCPwr_QueryMaxVoltageLevel (ViSession Vi,
                                        ViConstString ChannelName,
                                        ViReal64 CurrentLimit,
                                        ViReal64* MaxVoltageLevel);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The name of the output on which to query the maximum voltage level. | ViConstString |
| CurrentLimit | The current limit for which to determine the maximum programmable voltage level. | ViReal64 |

| Outputs | Description | Base Type |
| --- | --- | --- |
| Return value (.NET) | Returns the maximum programmable voltage level for the specified current limit. | ViReal64 |
| MaxVoltageLevel (C/COM) | Returns the maximum programmable voltage level for the specified current limit. | ViReal64* |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

### Query Output State

Description

This function returns whether the power supply is in a particular output state.

A constant voltage condition occurs when the output voltage is equal to the value of the Voltage Level attribute and the current is less than or equal to the value of the Current Limit attribute.

A constant current condition occurs when the output current is equal to the value of the Current Limit attribute and the Current Limit Behavior attribute is set to the Current Regulate defined value.

An unregulated condition occurs when the output voltage is less than the value of the Voltage Level attribute and the current is less than the value of the Current Limit attribute.

An over-voltage condition occurs when the output voltage is equal to or greater than the value of the OVP Limit attribute and the OVP Enabled attribute is set to True.

An over-current condition occurs when the output current is equal to or greater than the value of the Current Limit attribute and the Current Limit Behavior attribute is set to the Current Trip defined value.

When either an over-voltage condition or an over-current condition occurs, the power supply‘s output protection disables the output.  If the power supply is in an over-voltage or over-current state, it does not produce power until the output protection is reset.  The Reset Output Protection function resets the output protection.  Once the output protection is reset, the power supply resumes generating a power signal.

.NET Method Prototype

Boolean Outputs[].QueryState(OutputState outputState);

COM Method Prototype

HRESULT Outputs.Item().QueryState([in] IviDCPwrOutputStateEnum OutputState,
       [out, retval] VARIANT_BOOL *InState);

C Prototype

ViStatus IviDCPwr_QueryOutputState (ViSession  Vi,
                                    ViConstString ChannelName,
                                    ViInt32 OutputState,
                                    ViBoolean* InState);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The output on which to query the output state. | ViConstString |
| OutputState | The output state for which to query. | ViInt32 |

| Outputs | Description | Base Type |
| --- | --- | --- |
| InState (.NET) | Returns True if the power supply is currently in the state specified with the OutputState parameter, and False if it is not. | ViBoolean |
| InState (C/COM) | Returns True if the power supply is currently in the state specified with the OutputState parameter, and False if it is not. | ViBoolean* |

Values for OutputState

| Name | Description |  |  |  |
| --- | --- | --- | --- | --- |
|  |  | Language |  | Identifier |
| Constant Voltage | Return value indicates whether a constant voltage condition exists. |  |  |  |
|  |  | C | IVIDCPWR_VAL_OUTPUT_CONSTANT_VOLTAGE |  |
|  |  | .NET | OutputState.ConstantVoltage |  |
|  |  | COM | IviDCPwrOutputConstantVoltage |  |
| Constant Current | Return value indicates whether a constant current condition exists. |  |  |  |
|  |  | C | IVIDCPWR_VAL_OUTPUT_CONSTANT_CURRENT |  |
|  |  | .NET | OutputState.ConstantCurrent |  |
|  |  | COM | IviDCPwrOutputConstantCurrent |  |
| Over Voltage | Return value indicates whether an over-voltage condition exists. |  |  |  |
|  |  | C | IVIDCPWR_VAL_OUTPUT_OVER_VOLTAGE |  |
|  |  | .NET | OutputState.OverVoltage |  |
|  |  | COM | IviDCPwrOutputOverVoltage |  |
| Over Current | Return value indicates whether an over-current condition exists. |  |  |  |
|  |  | C | IVIDCPWR_VAL_OUTPUT_OVER_CURRENT |  |
|  |  | .NET | OutputState.OverCurrent |  |
|  |  | COM | IviDCPwrOutputOverCurrent |  |
| Unregulated | Return value indicates whether an unregulated condition exists. |  |  |  |
|  |  | C | IVIDCPWR_VAL_OUTPUT_UNREGULATED |  |
|  |  | .NET | OutputState.Unregulated |  |
|  |  | COM | IviDCPwrOutputUnregulated |  |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

Compliance Notes

- If an IVI-C class driver defines additional values for the OutputState parameter, the actual values shall be greater than or equal to IVIDCPWR_VAL_OUTPUT_STATE_CLASS_EXT_BASE and less than IVIDCPWR_VAL_OUTPUT_STATE_SPECIFIC_EXT_BASE.

- If an IVI-C specific driver defines additional values for the OutputState parameter, the actual values shall be greater than or equal to IVIDCPWR_VAL_OUTPUT_STATE_SPECIFIC_EXT_BASE.

### Reset Output Protection

Description

This function resets the power supply output protection after an over-voltage or over-current condition occurs.

An over-voltage condition occurs when the output voltage is equal to or greater than the value of the OVP Limit attribute and the OVP Enabled attribute is set to True.

An over-current condition occurs when the output current is equal to or greater than the value of the Current Limit attribute and the Current Limit Behavior attribute is set to Current Trip.

When either an over-voltage condition or an over-current condition occurs, the output protection of the power supply disables the output. Once the output protection is reset, the power supply resumes generating a power signal.

Use the Query Output State function to determine if the power supply is in an over-voltage or over-current state.

.NET Method Prototype

void Outputs[].ResetOutputProtection();

COM Method Prototype

HRESULT Outputs.Item().ResetOutputProtection();

C Prototype

ViStatus IviDCPwr_ResetOutputProtection (ViSession Vi,
                                         ViConstString ChannelName);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The output on which to reset output protection | ViConstString |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

## IviDCPwrBase Behavior Model

After the user calls the Initialize or Reset functions, the power supply produces a power signal based on its current configuration.

All changes to the power supply’s IviDCPwrBase functions and attributes take place immediately.

# IviDCPwrTrigger Extension Group

## IviDCPwrTrigger Overview

The IviDCPwrTrigger extension group defines extensions for DC power supplies capable of changing the output signal based on a trigger event.

## IviDCPwrTrigger Attributes

The IviDCPwrTrigger capability group defines the following attributes:

- Trigger Source

- Triggered Current Limit

- Triggered Voltage Level

This section describes the behavior and requirements of each attribute.  The actual value for each attribute ID is defined in section 8, IVIDCPwr Attribute ID Definitions.

### Trigger Source

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViInt32 (C/COM) | R/W | OutputChannel | None | Configure Trigger Source |
| ViString (.NET) | R/W | OutputChannel | None | Configure Trigger Source |

.NET Property Name

Outputs[].TriggerSource

COM Property Name

Outputs.Item().TriggerSource

COM Enumeration Name

IviDCPwrTriggerSourceEnum

C Constant Name

IVIDCPWR_ATTR_TRIGGER_SOURCE

Description

Specifies the trigger source.  After an Initiate call, the power supply waits for a trigger event from the source specified with this attribute.  After a trigger event occurs, the power supply changes the voltage level to the value of the Triggered Voltage Level attribute and the current limit to the value of the Triggered Current Limit attribute.

Defined Values

In IVI.NET the trigger source is a string.  If an IVI driver supports a trigger source and the trigger source is listed in IVI-3.3 Cross Class Capabilities Specification, Section 3, then the IVI driver shall accept the standard string for that trigger source.  This attribute is case insensitive, but case preserving.  That is, the setting is case insensitive but when reading it back the programmed case is returned.  IVI specific drivers may define new trigger source strings for trigger sources that are not defined by IVI-3.3 Cross Class Capabilities Specification if needed.

| Name | Description |  |  |  |
| --- | --- | --- | --- | --- |
|  |  | Language |  | Identifier |
| Immediate | The power supply does not wait for a trigger before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_IMMEDIATE |  |
|  |  | COM | IviDCPwrTriggerSourceImmediate |  |
| External | The power supply waits for an external trigger before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_EXTERNAL |  |
|  |  | COM | IviDCPwrTriggerSourceExternal |  |
| Software Trigger | The power supply waits for the Send Software Trigger function to execute before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_SOFTWARE_TRIG |  |
|  |  | COM | IviDCPwrTriggerSourceSwTrigFunc |  |
| TTL0 | The power supply waits for a trigger on the TTL 0 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_TTL0 |  |
|  |  | COM | IviDCPwrTriggerSourceTTL0 |  |
| TTL1 | The power supply waits for a trigger on the TTL 1 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_TTL1 |  |
|  |  | COM | IviDCPwrTriggerSourceTTL1 |  |
| TTL2 | The power supply waits for a trigger on the TTL 2 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_TTL2 |  |
|  |  | COM | IviDCPwrTriggerSourceTTL2 |  |
| TTL3 | The power supply waits for a trigger on the TTL 3 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_TTL3 |  |
|  |  | COM | IviDCPwrTriggerSourceTTL3 |  |
| TTL4 | The power supply waits for a trigger on the TTL 4 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_TTL4 |  |
|  |  | COM | IviDCPwrTriggerSourceTTL4 |  |
| TTL5 | The power supply waits for a trigger on the TTL 5 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_TTL5 |  |
|  |  | COM | IviDCPwrTriggerSourceTTL5 |  |
| TTL6 | The power supply waits for a trigger on the TTL 6 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_TTL6 |  |
|  |  | COM | IviDCPwrTriggerSourceTTL6 |  |
| TTL7 | The power supply waits for a trigger on the TTL 7 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_TTL7 |  |
|  |  | COM | IviDCPwrTriggerSourceTTL7 |  |
| ECL0 | The power supply waits for a trigger on the ECL 0 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_ECL0 |  |
|  |  | COM | IviDCPwrTriggerSourceECL0 |  |
| ECL1 | The power supply waits for a trigger on the ECL 1 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_ECL1 |  |
|  |  | COM | IviDCPwrTriggerSourceECL1 |  |
| PXI Star | The power supply waits for a trigger on the PXI Star line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_PXI_STAR |  |
|  |  | COM | IviDCPwrTriggerSourcePXIStar |  |
| RTSI0 | The power supply waits for a trigger on the RTSI 0 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_RTSI_0 |  |
|  |  | COM | IviDCPwrTriggerSourceRTSI0 |  |
| RTSI1 | The power supply waits for a trigger on the RTSI 1 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_RTSI_1 |  |
|  |  | COM | IviDCPwrTriggerSourceRTSI1 |  |
| RTSI2 | The power supply waits for a trigger on the RTSI 2 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_RTSI_2 |  |
|  |  | COM | IviDCPwrTriggerSourceRTSI2 |  |
| RTSI3 | The power supply waits for a trigger on the RTSI 3 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_RTSI_3 |  |
|  |  | COM | IviDCPwrTriggerSourceRTSI3 |  |
| RTSI4 | The power supply waits for a trigger on the RTSI 4 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_RTSI_4 |  |
|  |  | COM | IviDCPwrTriggerSourceRTSI4 |  |
| RTSI5 | The power supply waits for a trigger on the RTSI 5 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_RTSI_5 |  |
|  |  | COM | IviDCPwrTriggerSourceRTSI5 |  |
| RTSI6 | The power supply waits for a trigger on the RTSI 6 line before changing the output signal. |  |  |  |
|  |  | C | IVIDCPWR_VAL_TRIG_RTSI_6 |  |
|  |  | COM | IviDCPwrTriggerSourceRTSI6 |  |

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

Compliance Notes

- If an IVI-C class driver defines additional values for this attribute, the actual values shall be greater than or equal to IVIDCPWR_VAL_TRIG_SRC_CLASS_EXT_BASE and less than IVIDCPWR_VAL_TRIG_SRC_SPECIFIC_EXT_BASE.

- If an IVI-C specific driver defines additional values for this attribute, the actual values must be greater than or equal to IVIDCPWR_VAL_TRIG_SRC_SPECIFIC_EXT_BASE.

- If a specific driver implements any of the defined values in the following table, it must also implement the corresponding capability group:

| Value | Required Capability Group |
| --- | --- |
| Software Trigger | IviDCPwrSoftwareTrigger |

- When an IVI-COM specific driver implements this attribute with additional elements in its instrument specific interfaces, the actual values of the additional elements shall be greater than or equal to Trigger Source Specific Ext Base.

### Triggered Current Limit

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 | R/W | OutputChannel | None | Configure Triggered Current Limit |

.NET Property Name

Outputs[].TriggeredCurrentLimit

COM Property Name

Outputs.Item().TriggeredCurrentLimit

C Constant Name

IVIDCPWR_ATTR_TRIGGERED_CURRENT_LIMIT

Description

Specifies the value to which the power supply sets the current limit after a trigger event occurs.  The units are Amps.

After an Initiate call, the power supply waits for a trigger event from the source specified with the Trigger Source attribute. After a trigger event occurs, the power supply sets the current limit to the value of this attribute.

After a trigger occurs, the value of the Current Limit attribute reflects the new value to which the current limit has been set.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

### Triggered Voltage Level

| Data Type | Access | Applies to | Coercion | High Level Functions |
| --- | --- | --- | --- | --- |
| ViReal64 | R/W | OutputChannel | None | Configure Triggered Voltage Level |

.NET Property Name

Outputs[].TriggeredVoltageLevel

COM Property Name

Outputs.Item().TriggeredVoltageLevel

C Constant Name

IVIDCPWR_ATTR_TRIGGERED_VOLTAGE_LEVEL

Description

Specifies the value to  which the power supply sets the voltage level after a trigger event occurs. The units are Volts.

After an Initiate call, the power supply waits for a trigger event from the source specified with the Trigger Source attribute.  After a trigger event occurs, the power supply sets the voltage level to the value of this attribute.

After a trigger occurs, the value of the Voltage Level attribute reflects the new value to which the voltage level has been set.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this property.

## IviDCPwrTrigger Functions

The IviDCPwrTrigger extension defines the following functions:

- Abort

- Configure Triggered Current Limit (IVI-C only)

- Configure Triggered Voltage Level (IVI-C only)

- Configure Trigger Source (IVI-C only)

- Initiate

This section describes the behavior and requirements of the function.

### Abort

Description

If the power supply is currently waiting for a trigger to change the output signal, this function returns the power supply to the ignore triggers state.

If the power supply is not waiting for a trigger, this function does nothing and returns Success.

.NET Method Prototype

void Trigger.Abort();

COM Method Prototype

HRESULT Trigger.Abort();

C Prototype

ViStatus IviDCPwr_Abort (ViSession Vi);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

### Configure Triggered Current Limit (IVI-C only)

Description

Configures the current limit the power supply attempts to generate after it receives a trigger.

.NET Method Prototype

N/A
(use the Outputs[].TriggeredCurrentLimit property)

COM Method Prototype

N/A

(use the Outputs.Item().TriggeredCurrentLimit property)

C Prototype

ViStatus IviDCPwr_ConfigureTriggeredCurrentLimit (ViSession Vi,
                                         ViConstString ChannelName,
                                         ViReal64 Limit);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The name of the output on which to configure the triggered current limit. | ViConstString |
| Limit | Specifies the current limit the power supply attempts to generate after it receives a trigger. The driver uses this value to set the Triggered Current Limit attribute. See the attribute description for more details. | ViReal64 |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

### Configure Triggered Voltage Level (IVI-C only)

Description

Configures the voltage level the power supply attempts to generate after it receives a trigger.

.NET Method Prototype

N/A

(use the Outputs[].TriggeredVoltageLevel property)

COM Method Prototype

N/A

(use the Outputs.Item().TriggeredVoltageLevel property)

C Prototype

ViStatus IviDCPwr_ConfigureTriggeredVoltageLevel (ViSession Vi,
                                         ViConstString ChannelName,
                                         ViReal64 Level);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The name of the output on which to configure the triggered voltage level. | ViConstString |
| Level | Specifies the voltage level the power supply attempts to generate after it receives a trigger. The driver uses this value to set the Triggered Voltage Level attribute. See the attribute description for more details. | ViReal64 |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

### Configure Trigger Source (IVI-C only)

Description

Configures the trigger source the power supply waits for before changing the output signal.

.NET Method Prototype

N/A

(use the Trigger.Source property)

COM Method Prototype

N/A

(use the Trigger.Source property)

C Prototype

ViStatus IviDCPwr_ConfigureTriggerSource (ViSession Vi,
                                          ViConstString ChannelName,
                                          ViInt32 Source);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The name of the output on which to configure the trigger source. | ViConstString |
| Source | Specifies the trigger source the power supply waits for before changing the output signal. The driver uses this value to set the Trigger Source attribute. See the attribute description for more details. | ViInt32 |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

### Initiate

Description

If the power supply is not currently waiting for a trigger, this function causes the power supply to wait for a trigger.

If the power supply is already waiting for a trigger, this function does nothing and returns Success.

.NET Method Prototype

void Trigger.Initiate();

COM Method Prototype

HRESULT Trigger.Initiate();

C Prototype

ViStatus IviDCPwr_Initiate (ViSession Vi);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

## IviDCPwrTrigger Behavior Model

The following behavior model shows the relationship between the IviDCPwrTrigger capability group and power supply behavior.

- Figure 5-2.  IviDCPwrTrigger Behavior Model

After the user calls the Initialize or Reset functions, the power supply enters the ignore trigger state.

In the ignore trigger state, the attributes of the IviDCPwrBase capability group determine the signal that the power supply produces.  IviDCPwrTrigger attributes can be set, but do not affect the output signal.

Calling Initiate moves the power supply to the wait for trigger state.

In the wait for trigger state, the power supply waits for a trigger on the trigger source specified by the Trigger Source attribute.  After the power supply receives a trigger, it sets the voltage level to the value of the Triggered Voltage Level attribute, and the current limit to the value of the Triggered Current Limit attribute.  It then returns to the ignore trigger state.

After the changes in output occur, the Voltage Level and Current Limit attributes reflect the power supply’s new configuration..

Calling the Abort function moves the power supply from its current state to the ignore trigger state. If the power supply has not yet responded to a trigger, no change occurs to the voltage level or current limit.

# IviDCPwrSoftwareTrigger Extension Group

## IviDCPwrSoftwareTrigger Overview

The IviDCPwrSoftwareTrigger extension group defines extensions for DC power supplies capable of changing the output signal based on a software trigger event.

## IviDCPwrTrigger Functions

The IviDCPwrTrigger extension defines the following functions:

- Send Software Trigger

This section describes the behavior and requirements of this function.

### Send Software Trigger

Refer to IVI-3.3: Standard Cross Class Capabilities Specification, Section 2 Software Triggering Capability for the prototype and complete description of this function.

## IviDCPwrSoftwareTrigger Behavior Model

The IviDCPwrSoftwareTrigger Extension Group follows the behavior model of the IviDCPwrTrigger capability group. The only modification to the behavior model from the IviDCPwrTrigger capability group is the ability to send software triggers.

## IviDCPwrSoftwareTrigger Compliance Notes

- If an instrument driver implements the IviDCPwrSoftwareTrigger Extension Group, it shall implement the IviDCPwrTrigger Extension Group.

- If an instrument driver implements the IviDCPwrSoftwareTrigger Extension Group, it shall implement the Software Trigger defined value for the Trigger Source attribute.

# IviDCPwrMeasurement Extension Group

## IviDCPwrMeasurement Overview

The IviDCPwrMeasurement extension group defines extensions for DC power supplies capable of calculating various measurements such as voltage and current from the output signal.

## IviDCPwrMeasurement Functions

The IviDCPwrMeasurement extension defines the following function:

- Measure

This section describes the behavior and requirements of the function.

### Measure

Description

Takes a measurement on the output signal and returns the measured value.

.NET Method Prototype

Double Outputs[].Measure(Ivi.DCPwr.MeasurementType measurementType);

COM Method Prototype

HRESULT Outputs.Item().Measure(
[in] IviDCPwrMeasurementTypeEnum MeasurementType,
[out, retval] DOUBLE *Measurement);

C Prototype

ViStatus IviDCPwr_Measure (ViSession  Vi,
                           ViConstString ChannelName,
                           ViInt32 MeasurementType,
                           ViReal64* Measurement);

Parameters

| Inputs | Description | Base Type |
| --- | --- | --- |
| Vi | Instrument handle | ViSession |
| ChannelName | The output on which to take a measurement | ViConstString |
| MeasurementType | The type of measurement to take. | ViInt32 |

| Outputs | Description | Base Type |
| --- | --- | --- |
| Return value (.NET) | The measured value. | ViReal64 |
| Measurement (C/COM) | The measured value. | ViReal64* |

Defined Values for MeasurementType Parameter

| Name | Description |  |  |  |
| --- | --- | --- | --- | --- |
|  |  | Language |  | Identifier |
| Voltage | Voltage is measured. |  |  |  |
|  |  | C | IVIDCPWR_VAL_MEASURE_VOLTAGE |  |
|  |  | .NET | Ivi.DCPwr.MeasurementType.Voltage |  |
|  |  | COM | IviDCPwrMeasurementVoltage |  |
| Current | Current is measured. |  |  |  |
|  |  | C | IVIDCPWR_VAL_MEASURE_CURRENT |  |
|  |  | .NET | Ivi.DCPwr.MeasurementType.Current |  |
|  |  | COM | IviDCPwrMeasurementCurrent |  |

Return Values (C/COM)

The IVI-3.2:  Inherent Capabilities Specification defines general status codes that this function can return.

.NET Exceptions

The IVI-3.2: Inherent Capabilities Specification defines general exceptions that may be thrown, and warning events that may be raised, by this method.

Compliance Notes

- If an IVI-C class driver defines additional values for the measurementType parameter, the actual values shall be greater than or equal to IVIDCPWR_VAL_MEASURE_CLASS_EXT_BASE and less than IVIDCPWR_VAL_MEASURE_SPECIFIC_EXT_BASE.

- If an IVI-C specific driver defines additional values for the rangeType parameter, the actual values shall be greater than or equal to IVIDCPWR_VAL_MEASURE_SPECIFIC_EXT_BASE.

## IviDCPwrMeasurement Behavior Model

The IviDCPwrMeasurement Extension Group follows the behavior model of the IviDCPwrBase capability group. The only modification to the behavior model from the IviDCPwrBase capability group is the ability to take measurement on the output signal.

# IVIDCPwr Attribute ID Definitions

The following table defines the ID value for all IviDCPwr class attributes.

| Table 8-2. IviDCPwr Attributes ID Values |  |
| --- | --- |
| Attribute Name | ID Definition |
| IVIDCPWR_ATTR_VOLTAGE_LEVEL | IVI_CLASS_ATTR_BASE + 1 |
| IVIDCPWR_ATTR_OVP_ENABLED | IVI_CLASS_ATTR_BASE + 2 |
| IVIDCPWR_ATTR_OVP_LIMIT | IVI_CLASS_ATTR_BASE + 3 |
| IVIDCPWR_ATTR_CURRENT_LIMIT_BEHAVIOR | IVI_CLASS_ATTR_BASE + 4 |
| IVIDCPWR_ATTR_CURRENT_LIMIT | IVI_CLASS_ATTR_BASE + 5 |
| IVIDCPWR_ATTR_OUTPUT_ENABLED | IVI_CLASS_ATTR_BASE + 6 |
| IVIDCPWR_ATTR_CHANNEL_COUNT | IVI_CLASS_ATTR_BASE + 203 |
| IVIDCPWR_ATTR_TRIGGER_SOURCE | IVI_CLASS_ATTR_BASE + 101 |
| IVIDCPWR_ATTR_TRIGGERED_CURRENT_LIMIT | IVI_CLASS_ATTR_BASE + 102 |
| IVIDCPWR_ATTR_TRIGGERED_VOLTAGE_LEVEL | IVI_CLASS_ATTR_BASE + 103 |

In revision 1.0 of IVI 7: IviDCPwr Class Specification, the base attribute ID value was listed as: IVI_CLASS_PUBLIC_ATTR_BASE.  The constant ID definition was renamed to IVI_CLASS_ATTR_BASE.  The value of IVI_CLASS_ATTR_BASE is defined in IVI-3.1: Driver Architecture Specification.

# IVIDCPwr Attribute Value Definitions

This section specifies the actual value for each defined attribute value.

Current Limit Behavior

| Value Name | Language | Identifier | Actual Value |
| --- | --- | --- | --- |
| Current Regulate | C | IVIDCPWR_VAL_CURRENT_REGULATE | 0 |
|  | COM | IviDCPwrCurrentLimitRegulate | 0 |
|  | .NET | CurrentLimitBehavior.Regulate | 0 |
| Current Trip | C | IVIDCPWR_VAL_CURRENT_TRIP | 1 |
|  | COM | IviDCPwrCurrentLimitTrip | 1 |
|  | .NET | CurrentLimitBehavior.Trip | 1 |
| Current Limit Behavior Class Ext Base | C | IVIDCPWR_VAL_CURRENT_LIMIT_BEHAVIOR_CLASS_EXT_BASE | 500 |
|  | COM |  |  |
| Current Limit Behavior Specific Ext Base | C | IVIDCPWR_VAL_CURRENT_LIMIT_BEHAVIOR_SPECIFIC_EXT_BASE | 1000 |
|  | COM |  | 1000 |

Trigger Source

| Value Name | Language | Identifier | Actual Value |
| --- | --- | --- | --- |
| Immediate | C | IVIDCPWR_VAL_TRIG_IMMEDIATE | 0 |
|  | COM | IviDCPwrTriggerSourceImmediate | 0 |
| External | C | IVIDCPWR_VAL_TRIG_EXTERNAL | 1 |
|  | COM | IviDCPwrTriggerSourceExternal | 1 |
| Software Trigger | C | IVIDCPWR_VAL_SOFTWARE_TRIG | 2 |
|  | COM | IviDCPwrTriggerSourceSwTrigFunc | 2 |
| TTL0 | C | IVIDCPWR_VAL_TRIG_TTL0 | 3 |
|  | COM | IviDCPwrTriggerSourceTTL0 | 3 |
| TTL1 | C | IVIDCPWR_VAL_TRIG_TTL1 | 4 |
|  | COM | IviDCPwrTriggerSourceTTL1 | 4 |
| TTL2 | C | IVIDCPWR_VAL_TRIG_TTL2 | 5 |
|  | COM | IviDCPwrTriggerSourceTTL2 | 5 |
| TTL3 | C | IVIDCPWR_VAL_TRIG_TTL3 | 6 |
|  | COM | IviDCPwrTriggerSourceTTL3 | 6 |
| TTL4 | C | IVIDCPWR_VAL_TRIG_TTL4 | 7 |
|  | COM | IviDCPwrTriggerSourceTTL4 | 7 |
| TTL5 | C | IVIDCPWR_VAL_TRIG_TTL5 | 8 |
|  | COM | IviDCPwrTriggerSourceTTL5 | 8 |
| TTL6 | C | IVIDCPWR_VAL_TRIG_TTL6 | 9 |
|  | COM | IviDCPwrTriggerSourceTTL6 | 9 |
| TTL7 | C | IVIDCPWR_VAL_TRIG_TTL7 | 10 |
|  | COM | IviDCPwrTriggerSourceTTL7 | 10 |
| ECL0 | C | IVIDCPWR_VAL_TRIG_ECL0 | 11 |
|  | COM | IviDCPwrTriggerSourceECL0 | 11 |
| ECL1 | C | IVIDCPWR_VAL_TRIG_ECL1 | 12 |
|  | COM | IviDCPwrTriggerSourceECL1 | 12 |
| PXI Star | C | IVIDCPWR_VAL_TRIG_PXI_STAR | 13 |
|  | COM | IviDCPwrTriggerSourcePXIStar | 13 |
| RTSI0 | C | IVIDCPWR_VAL_TRIG_RTSI_0 | 14 |
|  | COM | IviDCPwrTriggerSourceRTSI0 | 14 |
| RTSI1 | C | IVIDCPWR_VAL_TRIG_RTSI_1 | 15 |
|  | COM | IviDCPwrTriggerSourceRTSI1 | 15 |
| RTSI2 | C | IVIDCPWR_VAL_TRIG_RTSI_2 | 16 |
|  | COM | IviDCPwrTriggerSourceRTSI2 | 16 |
| RTSI3 | C | IVIDCPWR_VAL_TRIG_RTSI_3 | 17 |
|  | COM | IviDCPwrTriggerSourceRTSI3 | 17 |
| RTSI4 | C | IVIDCPWR_VAL_TRIG_RTSI_4 | 18 |
|  | COM | IviDCPwrTriggerSourceRTSI4 | 18 |
| RTSI5 | C | IVIDCPWR_VAL_TRIG_RTSI_5 | 19 |
|  | COM | IviDCPwrTriggerSourceRTSI5 | 19 |
| RTSI6 | C | IVIDCPWR_VAL_TRIG_RTSI_6 | 20 |
|  | COM | IviDCPwrTriggerSourceRTSI6 | 20 |
| Trigger Source Class Ext Base | C | IVIDCPWR_VAL_TRIG_SRC_CLASS_EXT_BASE | 500 |
|  | COM |  |  |
| Trigger Source Specific Ext Base | C | IVIDCPWR_VAL_TRIG_SRC_SPECIFIC_EXT_BASE | 1000 |
|  | COM |  | 1000 |

# IviDCPwr Function Parameter Value Definitions

This section specifies the actual values for each function parameter that defines values.

Configure Output Range

| Parameter: RangeType |  |  |  |
| --- | --- | --- | --- |
| Value Name | Language | Identifier | Actual Value |
| Current | C | IVIDCPWR_VAL_RANGE_CURRENT | 0 |
|  | .NET | RangeType.Current | 0 |
|  | COM | IviDCPwrRangeCurrent | 0 |
| Voltage | C | IVIDCPWR_VAL_RANGE_VOLTAGE | 1 |
|  | .NET | RangeType.Voltage | 1 |
|  | COM | IviDCPwrRangeVoltage | 1 |
| Configure Output Range Class Ext Base | C | IVIDCPWR_VAL_RANGE_TYPE_CLASS_EXT_BASE | 500 |
|  | COM |  |  |
| Configure Output Range Specific Ext Base | C | IVIDCPWR_VAL_RANGE_TYPE_SPECIFIC_EXT_BASE | 1000 |
|  | COM |  |  |

Query Output State

| Parameter: OutputState |  |  |  |
| --- | --- | --- | --- |
| COM Enumeration Name: IviDCPwrOutputStateEnum |  |  |  |
| Value Name | Language | Identifier | Actual Value |
| Constant Voltage | C | IVIDCPWR_VAL_OUTPUT_CONSTANT_VOLTAGE | 0 |
|  | .NET | OuputState.ConstantVoltage | 0 |
|  | COM | IviDCPwrOutputConstantVoltage | 0 |
| Constant Current | C | IVIDCPWR_VAL_OUTPUT_CONSTANT_CURRENT | 1 |
|  | .NET | OuputState.ConstantCurrent | 1 |
|  | COM | IviDCPwrOutputConstantCurrent | 1 |
| Over Voltage | C | IVIDCPWR_VAL_OUTPUT_OVER_VOLTAGE | 2 |
|  | .NET | OuputState.OverVoltage | 2 |
|  | COM | IviDCPwrOutputOverVoltage | 2 |
| Over Current | C | IVIDCPWR_VAL_OUTPUT_OVER_CURRENT | 3 |
|  | .NET | OuputState.OverCurrent | 3 |
|  | COM | IviDCPwrOutputOverCurrent | 3 |
| Unregulated | C | IVIDCPWR_VAL_OUTPUT_UNREGULATED | 4 |
|  | .NET | OuputState.Unregulated | 4 |
|  | COM | IviDCPwrOutputUnregulated | 4 |
| Query Output State Class Ext Base | C | IVIDCPWR_VAL_OUTPUT_STATE_CLASS_EXT_BASE | 500 |
|  | COM |  |  |
| Query Output State Specific Ext Base | C | IVIDCPWR_VAL_OUTPUT_STATE_SPECIFIC_EXT_BASE | 1000 |
|  | COM |  |  |

Measure

| Parameter: MeasurementType |  |  |  |
| --- | --- | --- | --- |
| Value Name | Language | Identifier | Actual Value |
| Current | C | IVIDCPWR_VAL_MEASURE_CURRENT | 0 |
|  | .NET | MeasurementType.Current | 0 |
|  | COM | IviDCPwrMeasurementCurrent | 0 |
| Voltage | C | IVIDCPWR_VAL_MEASURE_VOLTAGE | 1 |
|  | .NET | MeasurementType.Voltage | 1 |
|  | COM | IviDCPwrMeasurementVoltage | 1 |
| Measure Class Ext Base | C | IVIDCPWR_VAL_MEASURE_CLASS_EXT_BASE | 500 |
|  | COM |  |  |
| Measure Specific Ext Base | C | IVIDCPWR_VAL_MEASURE_SPECIFIC_EXT_BASE | 1000 |
|  | COM |  |  |

# Error and Completion Code Value Definitions

The table below specifies the actual value for each status code, and the actual exception class and warning GUID for each IVI.NET exception and warning that the IviDCPwr class specification defines.  Note that warnings are raised as events in .NET.  See IVI-3.1 and IVI-3.2 for details.

| Table 11-1. IviDCPwr Error and Completion Codes |  |  |  |  |
| --- | --- | --- | --- | --- |
| Error Name | Description |  |  |  |
|  | API |  | Identifier | Value(hex) |
| Trigger Not Software | The trigger source is not set to software trigger. |  |  |  |
|  |  | .NET | Ivi.Driver. TriggerNotSoftwareException | IVI defined exception (see IVI 3.2) |
|  |  | C | IVIDCPWR_ERROR_TRIGGER_NOT_SOFTWARE | 0xBFFA1001 |
|  |  | COM | E_IVIDCPWR_TRIGGER_NOT_SOFTWARE | 0x80041001 |

# IviDCPwr Hierarchies

## IviDCPwr .NET Hierarchy

The full IviDCPwr .NET Hierarchy includes the Inherent Capabilities Hierarchy as defined in Section 4.1, .NET Inherent Capabilities of IVI-3.2: Inherent Capabilities Specification.  To avoid redundancy, the Inherent Capabilities are omitted here.

| Table 12-3. IviDCPwr .NET Hierarchy |  |  |
| --- | --- | --- |
| .NET Interface Hierarchy | Generic Name | Type |
| Outputs |  |  |
| Count | Output Channel Count | P |
| Name | Output Channel Name | P |
| [] | Item |  |
| ConfigureCurrentLimit | Configure Current Limit | M |
| ConfigureOvp | Configure OVP | M |
| ConfigureRange | Configure Output Range | M |
| CurrentLimit | Current Limit | P |
| CurrentLimitBehavior | Current Limit Behavior | P |
| Enabled | Output Enabled | P |
| Measure | Measure | M |
| OvpEnabled | OVP Enabled | P |
| OvpLimit | OVP Limit | P |
| QueryCurrentLimitMax | Query Current Limit Max | M |
| QueryState | Query Output State | M |
| QueryVoltageLevelMax | Query Voltage Level Max | M |
| ResetOutputProtection | Reset Output Protection | M |
| TriggeredCurrentLimit | Triggered Current Limit | P |
| TriggerSource | Trigger Source | P |
| TriggeredVoltageLevel | Triggered Voltage Level | P |
| VoltageLevel | Voltage Level | P |
| Trigger |  |  |
| Abort | Abort | M |
| Initiate | Initiate | M |
| SendSoftwareTrigger | Send Software Trigger | M |

### IviDCPwr .NET Interface Reference Properties

Interface reference properties are used to navigate the IviDCPwr .NET hierarchy. This section describes the interface reference properties that the IIviDCPwr interface defines.

| Data Type | .NET Property Name |
| --- | --- |
| IIviDCPwrOutputCollection | Outputs |
| IIviDCPwrOutput | Outputs[] |
| IIviDCPwrTrigger | Trigger |

## IviDCPwr COM Hierarchy

The full IviDCPwr COM Hierarchy includes the Inherent Capabilities Hierarchy as defined in Section 4.2, COM Inherent Capabilities of IVI-3.2: Inherent Capabilities Specification.  To avoid redundancy, the Inherent Capabilities are omitted here.

| Table 12-4. IviDCPwr COM Hierarchy |  |  |
| --- | --- | --- |
| COM Interface Hierarchy | Generic Name | Type |
| Outputs |  |  |
| Count | Output Channel Count | P |
| Name | Output Channel Name | P |
| Item |  |  |
| Enabled | Output Enabled | P |
| VoltageLevel | Voltage Level | P |
| TriggerSource | Trigger Source | P |
| TriggeredVoltageLevel | Triggered Voltage Level | P |
| TriggeredCurrentLimit | Triggered Current Limit | P |
| ConfigureCurrentLimit | Configure Current Limit | M |
| CurrentLimit | Current Limit | P |
| CurrentLimitBehavior | Current Limit Behavior | P |
| ConfigureOVP | Configure OVP | M |
| OVPLimit | OVP Limit | P |
| OVPEnabled | OVP Enabled | P |
| ConfigureRange | Configure Output Range | M |
| Measure | Measure | M |
| ResetOutputProtection | Reset Output Protection | M |
| QueryCurrentLimitMax | Query Current Limit Max | M |
| QueryVoltageLevelMax | Query Voltage Level Max | M |
| QueryState | Query Output State | M |
| Trigger |  |  |
| Abort | Abort | M |
| Initiate | Initiate | M |
| SendSoftwareTrigger | Send Software Trigger | M |

### IviDCPwr COM Interfaces

In addition to implementing IVI inherent capabilities interfaces, IviDCPwr interfaces contain interface reference properties for accessing the following IviDCPwr interfaces:

- IIviDCPwrOutputs

- IIviDCPwrTrigger

The IIviDCPwrOutputs interface contains methods and properties for accessing a collection of objects that implement the IIviDCPwrOutput interface.

Table 12-5. IviDCPwr Interface GUIDs lists the interfaces that this specification defines and their GUIDs.

| Table 12-5. IviDCPwr Interface GUIDs |  |
| --- | --- |
| Interface | GUID |
| IIviDCPwr | { 47ed51b6-a398-11d4-ba58-000064657374 } |
| IIviDCPwrOutputs | { 47ed51b7-a398-11d4-ba58-000064657374 } |
| IIviDCPwrOutput | { 47ed51b8-a398-11d4-ba58-000064657374 } |
| IIviDCPwrTrigger | { 47ed51b9-a398-11d4-ba58-000064657374 } |

### IviDCPwr COM Interface Reference Properties

Interface reference properties are used to navigate the IviDCPwr COM hierarchy. This section describes the interface reference properties that the IIviDCPwr interface defines.

| Data Type | COM Property Name |
| --- | --- |
| IIviDCPwrOutputs | Outputs |
| IIviDCPwrTrigger | Trigger |

### IviDCPwr COM Category

The IviDCPwr class COM Category shall be “IviDCPwr”, and the Category ID (CATID) shall be {47ed5153-a398-11d4-ba58-000064657374}.

## IviDCPwr C Function Hierarchy

The IviDCPwr class function hierarchy is shown in the following table. The full IviDCPwr C Function Hierarchy includes the Inherent Capabilities Hierarchy as defined in Section 4.3, C Inherent Capabilities of IVI-3.2: Inherent Capabilities Specification.  To avoid redundancy, the Inherent Capabilities are omitted here.

Table 12-6 IviDCPwr C Function Hierarchy

| Name or Class | Function Name |  |
| --- | --- | --- |
| Configuration |  |  |
| Configure Voltage Level | IviDCPwr_ConfigureVoltageLevel |  |
| Configure OVP | IviDCPwr_ConfigureOVP |  |
| Configure Current Limit | IviDCPwr_ConfigureCurrentLimit |  |
| Configure Output Range | IviDCPwr_ConfigureOutputRange |  |
| Configure Output Enabled | IviDCPwr_ConfigureOutputEnabled |  |
| Query Current Limit Max | IviDCPwr_QueryMaxCurrentLimit |  |
| Query Voltage Level Max | IviDCPwr_QueryMaxVoltageLevel |  |
| Triggering |  |  |
| Configure Trigger Source | IviDCPwr_ConfigureTriggerSource |  |
| Configure Triggered Voltage Level | IviDCPwr_ConfigureTriggeredVoltageLevel |  |
| Configure Triggered Current Limit | IviDCPwr_ConfigureTriggeredCurrentLimit |  |
| Action |  |  |
| Initiate | IviDCPwr_Initiate |  |
| Abort | IviDCPwr_Abort |  |
| Send SoftwareTrigger | IviDCPwr_SendSoftwareTrigger |  |
| Query Output State | IviDCPwr_QueryOutputState |  |
| Reset Output Protection | IviDCPwr_ResetOutputProtection |  |
| Measure | IviDCPwr_Measure |  |
| Get Output Channel Name | IviDCPwr_GetChannelName |  |

## IviDCPwr C Attribute Hierarchy

The IviDCPwr class attribute hierarchy is shown in the following table.  The full IviDCPwr C Attribute Hierarchy includes the Inherent Capabilities Hierarchy as defined in Section 4.3, C Inherent Capabilities of IVI-3.2: Inherent Capabilities Specification.  To avoid redundancy, the Inherent Capabilities are omitted here.

|  | Table 12-7. IviDCPwr C Attributes Hierarchy |  |
| --- | --- | --- |
|  | Category or Generic Attribute Name | C Defined Constant |
|  | Output |  |
|  | Current Limit | IVIDCPWR_ATTR_CURRENT_LIMIT |
|  | Current Limit Behavior | IVIDCPWR_ATTR_CURRENT_LIMIT_BEHAVIOR |
|  | Output Channel Count | IVIDCPWR_ATTR_CHANNEL_COUNT |
|  | Output Enabled | IVIDCPWR_ATTR_OUTPUT_ENABLED |
|  | OVP Enabled | IVIDCPWR_ATTR_OVP_ENABLED |
|  | OVP Limit | IVIDCPWR_ATTR_OVP_LIMIT |
|  | Voltage Level | IVIDCPWR_ATTR_VOLTAGE_LEVEL |
|  | Trigger |  |
|  | Trigger Source | IVIDCPWR_ATTR_TRIGGER_SOURCE |
|  | Triggered Current Limit | IVIDCPWR_ATTR_TRIGGERED_CURRENT_LIMIT |
|  | Triggered Voltage Level | IVIDCPWR_ATTR_TRIGGERED_VOLTAGE_LEVEL |

- Specific Driver Development Guidelines

  - Introduction

This section describes situations driver developers should be aware of when developing a specific instrument driver that complies with the IviDCPwr class.

  - Disabling Unused Extensions

Specific drivers are required to disable extension capability groups that an application program does not explicitly use. The specific driver can do so by setting the attributes of an extension capability group to the values that this section recommends. A specific driver can set these values for all extension capability groups when the Initialize or functions execute. This assumes that the extension capability groups remain disabled until the application program explicitly uses them. For the large majority of instruments, this assumption is true.

Under certain conditions, a specific driver might have to implement a more complex approach. For some instruments, configuring a capability group might affect instrument settings that correspond to an unused extension capability group. If these instrument settings affect the behavior of the instrument, then this might result in an interchangeability problem. If this can occur, the specific driver must take appropriate action so that the instrument settings that correspond to the unused extension capability group do not affect the behavior of the instrument when the application program performs an operation that might be affected by those settings.

The remainder of this section recommends attribute values that effectively disable each extension capability group.

- Disabling the IviDCPwrTrigger Extension Group

The IviDCPwrTrigger extension group affects the instrument behavior only when the user calls the Initiate function. Therefore, this specification does not recommend attribute values that disable the IviDCPwrTrigger extension group.

- Disabling the IviDCPwrSoftwareTrigger Extension Group

The IviDCPwrSoftwareTrigger extension group affects the instrument behavior only when the user calls the Send Software Trigger function. Therefore, this specification does not recommend attribute values that disable the IviDCPwrSoftwareTrigger extension group.

- Disabling the IviDCPwrMeasurement Extension Group

The IviDCPwrMeasurement extension group affects the instrument behavior only when the user calls the Measure function. Therefore, this specification does not recommend attribute values that disable the IviDCPwrMeasurement extension group.

  - Special Considerations for Configure Output Range

Some DC power supplies do not allow the user to explicitly specify an output’s range. Instead, they automatically change the range based on the values the user requests for the voltage level, OVP limit, and current limit. For instruments that automatically change the range, the IviDCPwr_ConfigureOutputRange function should perform range checking to verify that its input parameters are valid, but should not perform any communication with the instrument or set any attributes.

  - Special Considerations for Initiate

After an Initiate call, the values of the Voltage Level and Current Limit attributes might change. If a driver caches these attribute values, the driver needs to invalidate them when Initiate is called.

  - Special Considerations for Current Limit and Current Limit Behavior

Not all DC power supplies implement the current limit and current limit behavior settings in the same manner that IviDCPwr specification defines. Instead, some power supplies have a current level setting that is always active and an over-current limit setting that can be enabled or disabled. In order to develop a driver for these power supplies that complies with the IviDCPwr specification, the driver developer can implement the driver as described in this section.

It can be difficult to implement an IVI driver when the physical instrument settings differ from the attributes that the class specification defines. For these cases, a preferred approach can be to implement hidden attributes in the specific driver that control the physical instrument settings. The public attributes that the class specification defines manipulate the hidden attributes that correspond to the instrument settings. For this example the specific driver might implement the following hidden attributes

Instr Current Level – A hidden attribute that sets the instrument’s current level.

Instr OCP Limit – A hidden attribute that sets the instrument’s over-current protection limit.

Instr OCP Enable – A hidden attribute that enables and disables the instrument’s over-current protection.

Using these hidden attributes, the driver developer implements the specific driver as follows:

- When the user sets the Current Limit attribute, the specific driver sets both the Instr Current Level attribute and the Instr OCP Limit to the value that the user requests.

- When the user sets the Current Limit Behavior attribute to Current Trip, the driver sets the Instr OCP Enable to True. When the user sets the Current Limit Behavior attribute to Current Regulate, the driver sets the Instr OCP Enable to False.

For one particular DC power supply, the over-current protection is disabled by setting the over-current limit to a value that is 120% of the maximum current that the power supply can produce. In this case the specific driver implementation changes as follows:

- The specific driver implements hidden attributes only for the current level and the over-current limit.

- When the user sets the Current Limit attribute the specific driver always sets the Instr Current Level attribute to the value that the user requests. If the Current Limit Behavior attribute is set to Trip, the specific driver also sets the Instr OCP Limit to the value that the user requests. Otherwise, the specific driver does not set the value of the Instr OCP Limit attribute.

- When the user sets the Current Limit Behavior attribute to Current Regulate, the driver sets the Instr OCP Limit to 120% of the maximum current that the power supply can produce. When the user sets the Current Limit Behavior attribute to Current Trip, the driver sets the Instr OCP Limit to the same value as the Current Level attribute.

- Interchangeability Checking Rules

  - Introduction

IVI drivers have a feature called interchangeability checking. Interchangeability checking returns a warning when it encounters a situation were the application program might not produce the same behavior when the user attempts to use a different instrument.

  - When to Perform Interchangeability Checking

Interchangeability checking occurs when all of the following conditions are met:

- The Interchange Check attribute is set to True

- The user calls one of the following functions.

- Measure

- Initiate

  - Interchangeability Checking Rules

Interchangeability checking is performed on a capability group basis. When enabled, interchangeability checking is always performed on the base capability group. In addition, interchangeability checking is performed on extension capability groups for which the user has ever set any of the attributes of the group. If the user has never set any attributes of an extension capability group, interchangeability checking is not performed on that group.

In general interchangeability warnings are generated if the following conditions are encountered:

- An attribute that affects the behavior of the instrument is not in a state that the user specifies.

- The user sets a class driver defined attribute to an instrument-specific value.

- The user configures the value of an attribute that the class defines as read-only. In a few cases the class drivers define read-only attributes that specific drivers might implement as read/write.

The remainder of this section defines additional rules and exceptions for each capability group.

IviDCPwrBase Capability Group

No additional interchangeability rules or exceptions are defined for the IviDCPwrBase capability group.

IviDCPwrTrigger Capability Group

The driver needs to perform interchangeability checking on this group only when the function performing interchangeability checking is Initate.

IviDCPwrSoftwareTrigger Capability Group

No additional interchangeability rules or exceptions are defined for the IviDCPwrSoftwareTrigger capability group.

IviDCPwrMeasurement Capability Group

No additional interchangeability rules or exceptions are defined for the IviDCPwrMeasurement capability group.